Calorimetry Device

Initial Study

| Author | Checked | Approved |
| --- | --- | --- |
| Vasileios Gkionis | Name Surname | Name Surname |
| August 3 2026 | Date | Date |

Revision History

| Iss. Rev. | Date | Author | Comments/Change log |
| --- | --- | --- | --- |
| 1.0 | 03–08–2026 | G. Vasileios | Initial release |

Table of Contents

1	Introduction	4

2	Applicable And Reference Documents	4

2.1	Applicable documents	4

2.2	Reference Documents	4

3	Abbreviations And Acronyms	4

4	Topic	4

4.1	Subtopic	4

4.1.1	Subtopic	4

# Introduction

The purpose of this document is to report the full system-level study for the design, evaluation, and assembly of a calorimeter developed to directly measure the power loss of Roboteq's ultra-low-voltage (ULV) motor drives, specifically the F-series footprint (140 × 140 × 25 mm, conduction-plate cooled), rated at 1.5 kW.

At the efficiency levels these drives achieve, typically 90% to above 99%, the dissipated power can be as low as 15 W against a 1.5 kW throughput, rising to at most 200 W at the low end of that efficiency range. Measuring this loss as the difference between electrical input and output power is unreliable at this scale, since the quantity being measured is small relative to the uncertainty of the instrumentation used to measure the much larger quantities it is subtracted from. Calorimetry avoids this limitation by measuring the dissipated heat directly, rather than inferring it from two closely matched electrical measurements.

This document presents the trade study conducted across candidate calorimeter topologies, evaluated on accuracy, thermal response, build complexity, and cost, with particular attention to maintaining accuracy across the full 15–200 W loss envelope rather than at a single operating point. It then develops the thermal and control model of the selected configuration, a double-jacket chamber (DJC) with an actively servoed guard that nulls the temperature gradient across the inner wall, establishing the basis for meeting the ±1 W accuracy target and the hardware and instrumentation choices that follow from it.

# Applicable And Reference Documents

## Applicable documents

| Document ID | Document Title |
| --- | --- |

## Reference Documents

| Document ID | Document Title |
| --- | --- |

# Abbreviations And Acronyms

| NA | Not Applicable |
| --- | --- |
| DUT | Device Under Test |
| DJC | Double-Jacket Chamber |

# Calorimetry Theory and System Modeling

## Calorimeter

A calorimeter is a thermally isolated chamber in which a coolant loop provides, ideally, the only path for heat exchange between the chamber’s contents and the ambient room. The device under test (DUT) is placed inside the chamber and operated at its rated power. The power the DUT dissipates as loss heats the chamber and is carried away by the coolant loop. Because this heat has no other way out, it can be measured directly from the coolant’s flow rate and temperature rise.  This is what makes calorimetry well suited to high-precision loss measurement: it is agnostic to the source of the heat. A calorimeter cannot distinguish whether the dissipated power originated as conduction loss, switching loss, hysteresis loss, or a chemical reaction, it registers only the total heat crossing the boundary. The total loss of a DUT, whether a power converter, a motor, or any other heat-producing device, is therefore obtained directly from the heat it dissipates inside the chamber, rather than from a breakdown of individual loss mechanisms. This is why calorimetry is the primary method for characterizing the losses of power converters: these devices combine high power throughput with high efficiency, so the loss to be measured is a small quantity riding on a much larger one, a regime in which other measurement approaches lose precision.

## Topology evaluation

At the beginning of the study, 3 topologies were evaluated and compared for the most fitting candidate to be used as a template for this assignment. Those were a single wall layer, a double jacket copper layer and a complete double jacket layer respectively.

### Subtopic

#### Subtopic

Content