Calorimetry Device

Initial Study

| Author | Checked | Approved |
| --- | --- | --- |
| Vasileios Gkionis | Name Surname | Name Surname |
| August 3 2026 | Date | Date |

Revision History

| Iss. Rev. | Date | Author | Comments/Change log |
| --- | --- | --- | --- |
| 1.0 | 03–08–2026 | G. Vasileios | Initial release |
| 1.1 | 30–08–2026 | G. Vasileios | Sections 2–7 completed: reference documents, calorimetry theory, topology trade study and selection, design decisions, thermal model, state-space model and guard control design, Python simulation and results. Sections 8–9 (hardware, firmware) still to be written. |

Table of Contents

1	Introduction	4

2	Applicable And Reference Documents	4

2.1	Applicable documents	4

2.2	Reference Documents	4

3	Abbreviations And Acronyms	4

4	Topic	4

4.1	Subtopic	4

4.1.1	Subtopic	4

# Introduction

The purpose of this document is to report the full system-level study for the design, evaluation, and assembly of a calorimeter developed to directly measure the power loss of Roboteq's ultra-low-voltage (ULV) motor drives, specifically the F-series footprint (140 × 140 × 25 mm, conduction-plate cooled), rated at 1.5 kW.

At the efficiency levels these drives achieve, typically 90% to above 99%, the dissipated power can be as low as 15 W against a 1.5 kW throughput, rising to at most 200 W at the low end of that efficiency range. Measuring this loss as the difference between electrical input and output power is unreliable at this scale, since the quantity being measured is small relative to the uncertainty of the instrumentation used to measure the much larger quantities it is subtracted from. Calorimetry avoids this limitation by measuring the dissipated heat directly, rather than inferring it from two closely matched electrical measurements.

This document presents the trade study conducted across candidate calorimeter topologies, evaluated on accuracy, thermal response, build complexity, and cost, with particular attention to maintaining accuracy across the full 15–200 W loss envelope rather than at a single operating point. It then develops the thermal and control model of the selected configuration, a double-jacket chamber (DJC) with an actively servoed guard that nulls the temperature gradient across the inner wall, establishing the basis for meeting the ±1 W accuracy target and the hardware and instrumentation choices that follow from it.

# Applicable And Reference Documents

## Applicable documents

| Document ID | Document Title |
| --- | --- |
| AD-01 | IEC 60034-2-1 — Rotating electrical machines: standard methods for determining losses and efficiency from tests |
| AD-02 | IEC 61800-9-2 — Adjustable speed electrical power drive systems: ecodesign, energy efficiency indicators for power drive systems and motor starters |
| AD-03 | Nidec Automation / Roboteq — Ultra-Low-Voltage Drive Catalog, 17 Sep 2025 (F-series DUT class: 140 × 140 × 25 mm, 2 × 60 A / 1 × 120 A) |

## Reference Documents

| Document ID | Document Title |
| --- | --- |
| RD-01 | Christen, D., Badstübner, U., Biela, J., Kolar, J. W., “Calorimetric Power Loss Measurement for Highly Efficient Converters”, IPEC 2010, ETH Zürich Power Electronic Systems Laboratory. Built air-gap double-jacket calorimeter, ±0.4 W / 1 %. Primary architecture reference. |
| RD-02 | Simpson, N., Hopkins, A. N., “An Accurate and Flexible Calorimeter Topology for Power Electronic System Loss Measurement”, IEMDC 2017, University of Bristol. Surface-heated closed calorimeter, < 1 % over 20–100 W. Guard-actuator and commissioning reference. |
| RD-03 | Nair, D. G., “Calorimetric Power Loss Measurement of a Small Power Converter”, MSc thesis, Aalto University, 2013. Source of the double-jacket schematic and of the guard criterion ¦Ti − Te¦ < Pacc·Rth(in) (§3.4.2). Also the unguarded counter-example. |
| RD-04 | Kosonen, A., et al., “Calorimetric concept for measurement of power losses up to 2 kW in electric drives”, IET Electric Power Applications, 2013. Modularity across DUT ratings. |
| RD-05 | Malliband, P. D., Carter, B., Gordon, M., Warne, D. F., “Accurate measurement of induction motor losses using balance calorimeter”, Proc. IEE, vol. 138, no. 5, 1991. |
| RD-06 | Malliband, P. D., et al., “Design of a double-jacketed, closed type calorimeter for direct measurement of motor losses”, 7th Int. Conf. on Power Electronics and Variable Speed Drives, no. 456, 1998. Original DJC proposal. |
| RD-07 | Blaabjerg, F., Pedersen, J. K., Ritchie, E., “Calorimetric measuring systems for characterizing high frequency power losses in power electronic components and systems”, IEEE IAS Conf., vol. 2, pp. 1369–1376, 2002. First proposal of active air-gap temperature control. |
| RD-08 | Xiao, C., Chen, G., Odendaal, W. G. H., “Overview of Power Loss Measurement Techniques in Power Electronic Systems”, IEEE Trans. Industry Applications, pp. 657–664, 2007. |
| RD-09 | Jalilian, A., Gosbell, V. J., Perera, B. S. P., Cooper, P., “Double chamber calorimeter (DCC): a new approach to measure induction motor harmonic losses”, IEEE Trans. Energy Conversion, vol. 14, no. 3, 1999. The series calorimeter — distinct from the DJC. |
| RD-10 | “Overview of calorimetric systems used in loss determination of electric motors and drives”, IEEE, 2017. Source of the four-way open / closed / balanced / series classification. |
| RD-11 | Internal working notes, Obsidian vault Roboteq Internship/Calorimetry/ — Study, thermo modeling, Parameter Calculation, Biot check, System Modeling Final, BOM manual, Zephyr map. |
| RD-12 | Simulation source, ~/roboteq/python/ — calorimetry-ss.py (plant), calorimetry_sim.py (closed loop, probes, plots), params.yaml. |

# Abbreviations And Acronyms

| ADC | Analog-to-Digital Converter |
| --- | --- |
| AW | Anti-Windup |
| B1 | Measurement boundary — the surface enclosing the inner chamber contents |
| BOM | Bill of Materials |
| DAQ | Data Acquisition |
| DC | Direct Current — here also the zero-frequency gain of a transfer function |
| DJC | Double-Jacket Chamber |
| DUT | Device Under Test |
| EPP | Expanded Polypropylene |
| ESD | Electrostatic Discharge |
| FF | Feedforward |
| HX | Heat Exchanger |
| I²C | Inter-Integrated Circuit (two-wire serial bus) |
| IMC | Internal Model Control — the λ-tuning rule used for the guard PI |
| LTI | Linear Time-Invariant |
| MCU | Microcontroller Unit |
| NA | Not Applicable |
| PI / PID | Proportional-Integral / Proportional-Integral-Derivative controller |
| PWM | Pulse Width Modulation |
| RK4 | Fourth-order Runge-Kutta fixed-step integrator |
| RSS | Root Sum of Squares |
| RTD | Resistance Temperature Detector (Pt100 / Pt1000) |
| SISO | Single-Input Single-Output |
| TIM | Thermal Interface Material |
| ULV | Ultra-Low Voltage (≤ 60 V — the Roboteq / Nidec Automation drive class) |
| XPS | Extruded Polystyrene (rigid foam insulation) |
| ZOH | Zero-Order Hold |

# Calorimetry Theory and System Modeling

## Calorimeter

A calorimeter is a thermally isolated chamber in which a coolant loop provides, ideally, the only path for heat exchange between the chamber’s contents and the ambient room. The device under test (DUT) is placed inside the chamber and operated at its rated power. The power the DUT dissipates as loss heats the chamber and is carried away by the coolant loop. Because this heat has no other way out, it can be measured directly from the coolant’s flow rate and temperature rise.  This is what makes calorimetry well suited to high-precision loss measurement: it is agnostic to the source of the heat. A calorimeter cannot distinguish whether the dissipated power originated as conduction loss, switching loss, hysteresis loss, or a chemical reaction, it registers only the total heat crossing the boundary. The total loss of a DUT, whether a power converter, a motor, or any other heat-producing device, is therefore obtained directly from the heat it dissipates inside the chamber, rather than from a breakdown of individual loss mechanisms. This is why calorimetry is the primary method for characterizing the losses of power converters: these devices combine high power throughput with high efficiency, so the loss to be measured is a small quantity riding on a much larger one, a regime in which other measurement approaches lose precision.

## Why calorimetry, and not an electrical measurement

The direct way to find a converter's loss is to measure input and output power electrically and subtract. That method degrades exactly where this project needs it most, because the measurand is the small difference of two large, nearly equal numbers.

For the F-series DUT at 1.5 kW, the loss to be resolved is between 15 W (η = 99 %) and 200 W (η ≈ 88 %). Resolving the 15 W point to ±1 W by subtraction requires each of the two 1.5 kW measurements to be accurate to better than ±0.7 W, i.e. ±0.05 % of reading. Real bench instrumentation does not do this:

| Power-analyser class | Uncertainty per channel at 1.5 kW | Uncertainty on Pin − Pout (RSS) | As a fraction of a 15 W loss |
| --- | --- | --- | --- |
| ±0.05 % of reading | 0.75 W | 1.06 W | 7.1 % |
| ±0.1 % | 1.5 W | 2.12 W | 14.1 % |
| ±0.2 % | 3.0 W | 4.24 W | 28.3 % |
| ±0.5 % | 7.5 W | 10.6 W | 70.7 % |

Table 1. Why subtraction fails at the high-efficiency end. The uncertainty is fixed by the throughput, not by the loss, so it does not shrink as the DUT gets better. The same instrument that reads a 200 W loss to 2 % reads a 15 W loss to 14 % or worse.

Calorimetry does not have this structure. It measures the dissipated heat itself, so its uncertainty is set by the flow meter, the temperature-pair match and the guard residual — all of which are roughly constant in watts across the range, and none of which depend on the DUT's throughput, waveform, switching frequency, power factor or EMI. A calorimeter that is good to ±1 W is good to ±1 W at 15 W and at 200 W alike.

| Aspect | Electrical (input − output) | Calorimetric |
| --- | --- | --- |
| What is measured | Two large electrical powers, then their difference | The heat leaving the DUT, directly |
| Dominant error | Difference of two similar large numbers — error grows as a fraction of the loss as efficiency rises | Absolute — stays roughly constant in watts |
| Sensitive to | Waveform distortion, phase/timing error, sensor bandwidth, EMI | Ambient stability, wall leakage, calibration drift |
| Electrical connection to DUT | Required | None — fully decoupled |
| Speed | Seconds | Minutes to hours per point (thermal settling) |
| Best suited to | Low to moderate efficiency, transient measurement | Verifying loss on high-efficiency devices (η > ~95 %) |

Table 2. The two methods compared. Calorimetry is slower and needs more hardware; it buys accuracy that does not degrade with the DUT's efficiency.

Published, well-built calorimeters report about ±0.4 W absolute, or 0.2–1 % of full scale (RD-01, RD-02), essentially flat across the power range. The accuracy target adopted for this rig is Pacc = ±1 W, chosen as a realistic first-build figure with margin against the literature best. At 1.5 kW throughput, ±1 W resolves 0.07 efficiency points — enough to separate 97.0 % from 97.5 % (a 7.5 W step) without ambiguity.

## Calorimeter families

Calorimeters built for power-loss measurement are classified into four chamber types (RD-10). Each can be run air- or liquid-cooled.

| Type | How it works | Trade-off |
| --- | --- | --- |
| Open | Single pass of coolant drawn from ambient through the chamber; loss from inlet/outlet ΔT and flow | Simplest build; accuracy is at the mercy of ambient temperature and flow fluctuations |
| Closed | DUT sealed in an insulated enclosure with a closed coolant loop that dumps to an external exchanger | Much better ambient isolation; usually built with a guarded (double-jacket) wall to stop the residual leak |
| Balanced (substitution) | A trim heater holds the chamber at a fixed balance point; loss = the reduction in trim-heater power relative to a DUT-off baseline | Very accurate and needs no flow measurement, but requires a control loop and a baseline calibration run |
| Series | Two chambers share one coolant loop — one holds a reference heater of known power, the other the DUT — so the rig self-calibrates continuously | Best drift immunity, at the cost of a second full chamber |

Table 3. The four chamber families. This project builds a closed, guarded, liquid-cooled chamber with a flow readout, and keeps substitution as the calibration method rather than as the readout.

Two readout philosophies, kept separate A guarded chamber can report the loss in two different ways, and the project notes use both words. Flow readout (this build): the loss is ṁ·cp·ΔTw through the water loop, and the guard exists only to null the wall so that reading is complete. Substitution readout: the loss is a difference of trim-heater powers, and no flow measurement is needed. This rig uses the flow readout for the measurement and the substitution principle only for commissioning — the inner-chamber resistor injects a known electrical power that the flow readout must reproduce.

## Reference builds

Four published builds set the design envelope and supplied the specific techniques that were copied. Three are used as positive references; the fourth is the counter-example that justifies the guard.

| Source | Topology | DUT / loss range | Reported accuracy | What was taken from it |
| --- | --- | --- | --- | --- |
| ETH Zürich, IPEC 2010 (RD-01) | Air-gap double-jacket, actively servoed gap | 5 kW DC-DC at 99.1 % η; also a 1.6 kW PFC at 15 W loss | ±0.4 W or 1 %, roughly flat over 10–100 W | The architecture itself: nested shells, servoed air gap, closed water loop with its own radiator and fans, in-chamber preheat to shorten settling, digital I²C wall sensors |
| Bristol, IEMDC 2017 (RD-02) | Closed single box with a PI-controlled copper-tape surface heater — the guard collapsed onto the wall | 20–100 W (predicted ceiling ≈ 250 W) | < 1 % at all 12 load points; < 0.4 W absolute | The guard as one software PI loop; in-situ gravimetric flow calibration of a cheap flow sensor; automatic steady-state detection from the temperature slope; commodity PC-watercooling parts; single-point calibration |
| Kosonen et al., IET EPA 2013 (RD-04) | One rig covering several drive ratings | Up to 2 kW of loss | — | The modularity argument: build one chamber and re-range it per DUT class rather than one rig per product |
| Nair, Aalto MSc 2013 (RD-03) | Closed single-wall, unguarded chamber (a DJC was reviewed, then not built) | 0.75 kW converter at ~97 % η | ±1.5 % below 50 W, degrading with power | The counter-example. Its own conclusion names wall leakage as the dominant error. Also the source of the DJC schematic and of the guard criterion used in §4.6.3 |

Table 4. The four reference builds. Nair and ETH differ almost only in whether the wall gradient is actively nulled, and their accuracy figures differ by roughly a factor of four — which is the whole argument for the guard.

Two of these builds, a decade apart and unrelated, independently converged on the same commodity parts: XPS/EPP foam shells, Aqua Computer PC-watercooling radiators, a peristaltic pump, a cheap flow sensor rescued by gravimetric calibration, and Class-A platinum RTDs. That convergence is why this project's BOM starts from the same categories rather than from laboratory-instrument catalogues.

## Topology evaluation

At the beginning of the study, 3 topologies were evaluated and compared for the most fitting candidate to be used as a template for this assignment. Those were a single wall layer, a double jacket copper layer and a complete double jacket layer respectively.

Those three are compared in §4.5.1; the reasoning that led to them, starting from a wider four-way trade study, is in §4.5.2, and the decision in §4.5.3. The remainder of this section describes the selected topology and the design choices made during the study. Sections 5 to 7 then develop the model of that topology in increasing detail: the thermal network (§5), its state-space form and the control design that follows from it (§6), and the Python simulation used to size the actuators and predict the settling schedule (§7).

### The three candidate topologies

|  | 1 — Single wall, passive | 2 — Surface-heated single box | 3 — Full air-gap double jacket |
| --- | --- | --- | --- |
| Shells to build | 1 | 1 | 2, nested |
| Guard actuator | None — insulation only | Copper-foil strip heater bonded to the outer surface, PI-controlled | Heated and fanned air gap between the shells, PI-controlled |
| Guard sensing | n/a | Surface RTDs, ~1 mm under the skin | Gap-air temperature Te vs chamber air Ti |
| Feedthroughs cross | 1 wall | 1 wall | 2 walls |
| Wall leak | Slowed, never cancelled; drifts with ambient | Actively nulled at the wall surface | Actively nulled across the whole inner wall |
| Demonstrated accuracy | ±1.5 %, worsening with power (RD-03) | < 1 % over 20–100 W (RD-02) | ±0.4 W flat (RD-01) |
| Estimated cost | Lowest | ≈ €620–1 020 | ≈ €710–1 150 |
| Main risk | Accuracy — will not meet ±1 W | Heater and temperature uniformity over a small box | Build effort: a second shell to fabricate and seal |

Table 5. The three topologies. The measurement core — water loop, flow meter, ΔT pair, in-chamber radiator and fans — is identical in all three and costs about €490–750; the whole difference is the guard.

### How the shortlist was reached

The study did not start with these three. The first pass, before the literature review, compared four configurations against cost, build time, thermal time constant and accuracy:

|  | A — Open, low-cost | B — Closed, passive guard | C — Closed, active guard | D — Repurposed chamber |
| --- | --- | --- | --- | --- |
| Readout | Single-pass air, ṁ·cp·ΔT | Recirculating air in an insulated box, ṁ·cp·ΔT | Trim-heater power difference (substitution) | Same as B, retrofitted into an existing enclosure |
| Parts cost | ~$100–150 | ~$500–950 | ~$450–800 | ~$530–810 if a chamber is on hand; $1 800–3 300 otherwise |
| Accuracy at 150–200 W | ±10–20 % | ±3–5 % | ±1–2 % | ±3–7 % |
| Accuracy at 15–30 W | ±10–20 % | ±15–30 % | ±1–2 % | ±10–20 % |
| Build effort | ~1 week | ~2–3 weeks | ~4–6 weeks | ~1–2 weeks if a chamber exists |
| Outcome | Kept only as a week-1 shakedown of the sensor and logging chain | Fallback if schedule ran out | Selected | Ruled out on day 1 — no chamber exists at Roboteq HQ to repurpose |

Table 6. The original four-way trade study. Cost figures are 2026 order-of-magnitude estimates for individual components, not vendor quotes.

Two findings decided it. First, B's weakness is at the wrong end of the range: a flow sensor's and an RTD pair's uncertainties are fixed in absolute terms, so at the 200 W / 10 K design point they are a small fraction of the reading, but at 15 W and 0.75 K of air-side rise the same fixed ±0.15 K is 15–20 % of the signal. Relative accuracy therefore degrades exactly at the high-efficiency operating points calorimetry exists to serve. Second, once the DUT's real loss range was known, C was no longer the more expensive option — it removes the precision flow sensor that dominates B's cost and replaces it with foam, a resistor and a control loop. C dominated B on both cost and accuracy; the only axis where B still won was schedule risk.

The literature review then reshaped what "active guard" means in hardware. The Bristol paper (RD-02) showed the same guard physics implemented as a heater bonded to the outside of a single box, which removes the second shell — the main source of C's schedule risk. That turned the original A/B/C/D question into the three-way comparison of Table 5, all three columns of which come from real, published, measured builds rather than from estimates.

### Decision

Topology 3 — the full air-gap double-jacket chamber — is the build, confirmed 2026-07-10 and reaffirmed 2026-07-13. Topology 2 stays documented as the fallback if the air-gap guard underperforms. The reasoning:

•	A passive wall cannot meet ±1 W. Worked for this rig's own geometry: at a chamber 10 K above ambient and Rth(in) ≈ 1.8 K/W, an unguarded inner wall leaks about 5.6 W. Even a calibrated leak model leaves a residual from ambient drift, draught-driven film-coefficient variation and seam ageing — realistically ~10 % of the correction, i.e. 0.5–0.6 W, before any other error term is counted. Nair (RD-03) measured exactly this failure mode.

•	The ETH air-gap variant is the best-characterised precedent at our accuracy target and is the one whose schematic, criterion and preheat trick the design already follows.

•	Its cost premium over the surface-heated variant is small (≈ €90–130) and its main risk — fabricating and sealing a second shell — is largely a purchasing question, since insulated foam boxes are a commodity product that can be ordered close to size rather than built from sheet stock.

•	Uniformity risk favours the air gap on a box this size. The Bristol surface heater was validated on a 525 × 305 × 195 mm box with all faces driven in series, and per-face control is listed there as future work. On a smaller chamber, a stirred air gap distributes the guard's heat for us, which a bonded foil heater does not.

## Selected topology — the double-jacket chamber

### Architecture

Two nested insulated shells. The inner chamber holds the DUT, the water-loop pickup radiator, circulation fans, a preheat/calibration resistor and the chamber temperature sensor. The gap between the shells has its own heater and fan and is servoed to track the inner chamber. The water loop crosses both walls and is the only intended exit for the DUT's heat.

Ta — room / ambient, free to drift

┌────────────────────────────────────────────────────────────┐

│  Outer shell — insulated, heated and fanned                │

│  air gap servoed to track the inner chamber                │

│                                                            │

│        Te — gap air        Pheater(e), Pfan(e)             │

│   ┌────────────────────────────────────────────────┐       │

│   │  Inner chamber                    B1           │       │

│   │                                                │       │

│   │   water in  ──►[ pickup radiator ]──► water out│       │

│   │                        ▲                       │       │

│   │                     [ DUT ] ──► Ploss          │       │

│   │                                                │       │

│   │   Ti — chamber air     Pheater(i), Pfan(i)     │       │

│   └────────────────────────────────────────────────┘       │

│      inner wall:  Rth(in) = d /(λ·A) ,  Ggap               │

└────────────────────────────────────────────────────────────┘

outer wall:  Rth(out) ,  Gout        → to ambient

The labelling follows RD-03 Fig. 3.4 exactly: Ti inner chamber, Te gap air, Ta ambient, with a heater and a fan on each side of the inner wall. Three boundaries matter, and every equation in §5 is a first-law balance across one of them:

| Boundary | Encloses | What crosses it |
| --- | --- | --- |
| B1 | The inner-chamber contents — the measurement boundary | In: the DUT's loss (as electricity) and the metered auxiliaries. Out: the water's enthalpy rise (measured) and the wall leak (the error term the guard deletes). |
| B2 | The air gap — the guard's boundary | In: gap heater and gap fan power, plus whatever leaks in from B1 (ideally zero). Out: loss to ambient through the outer wall — unmetered, and harmless, because it is not on the measured path. |
| B3 | Any single lump inside, when its own temperature is needed | Used for the foam wall and the DUT body only. |

Table 7. The three boundaries.

### Working principle

1.	The DUT is mounted in the inner chamber and run at its operating point. All of its loss becomes heat inside B1.

2.	The guard loop drives the gap air to the chamber's own temperature, so the gradient across the inner wall goes to zero.

3.	With no gradient, Fourier's law gives no conduction through that wall — the leak is cancelled, not merely slowed. Radiation across the gap is nulled by the same condition.

4.	The only remaining exit for the heat is the water loop.

5.	The loss is read from the coolant's mass flow and its inlet-to-outlet temperature rise, with the metered in-chamber auxiliaries subtracted.

Ploss = ṁ · cp,w · ΔTw − Paux(i)		(1)

This is the whole instrument. Everything else in this report exists either to justify the two conditions in steps 2 and 3, or to price how imperfectly they are met.

### The guard criterion and the two gates

Equation (1) is only valid while two conditions hold. Both are checked in firmware before a point is logged, and both are stated as explicit gates.

Null gate. From RD-03 §3.4.2, the residual gradient allowed across the inner wall is fixed by the accuracy budget and the wall's own resistance:

| Ti − Te |  <  Pacc · Rth(in)		(2)

With Pacc = 1 W and the design wall (XPS, 50 mm, λ ≈ 0.035–0.04 W/m·K over ≈ 0.6 m² ⇒ Ggap ≈ 0.56 W/K, Rth(in) ≈ 1.79 K/W), the gate is |Ti − Te| < 1.79 K. This is a design spec, not a comparison against the unguarded case: decide how accurate the rig must be, and the wall's resistance converts that directly into how tightly the servo has to hold the null. A realistic servo residual of ±0.3 K costs 0.17 W.

Steady gate. The storage term must be dead, or the chamber's thermal mass is silently charging or discharging into the reading:

| dP̂meas/dt |  <  Pthr / τdom		(3)

The slope threshold is the residual power threshold divided by the slowest time constant the reading actually shows. It is computed per operating point rather than hard-coded, because τdom moves by a factor of about three across the power range. Equivalently, in temperature terms, holding a 0.2 W storage slice against Ci + Cw ≈ 2.45 kJ/K means |dT/dt| < 0.29 K/h.

Gate the reading, not the chamber temperature Ti is dominated by the fast air/water exchange mode and settles long before the reading does. A steady-state detector that watches Ti will go green while the water side is still transient. The gate must watch P̂meas itself.

A useful property of the architecture: the outer shell's leak to the room does not need to be controlled at all. The servo nulls one boundary only — chamber to gap. Whatever the gap then exchanges with the room never crosses the measured boundary, so those watts are accuracy-free. The room only has to be stable enough for the inner loop to track. This is the same principle as the guard ring in guarded-hot-plate conductivity testing (ASTM C177).

## Design decisions taken during the study

The design changed several times over the two months as specifications firmed up and as parts were confirmed. The table records what changed, when, and why, so that the working notes can be read in the right order.

| Decision | Date | Outcome | Reason |
| --- | --- | --- | --- |
| Repurpose an existing chamber | Day 1 | Ruled out | No chamber, incubator or environmental enclosure exists at Roboteq HQ. The rig is a ground-up build. |
| Coolant medium | Early July | Water, closed recirculating loop | Water's volumetric heat capacity is ~3 000–4 000× air's, so the same heat is carried by a manageable flow with a readable ΔT. Matches RD-01; RD-03's mains-fed open loop is explicitly what its own future-work section recommends moving away from. |
| Chamber topology | 2026-07-10, reaffirmed 07-13 | Air-gap DJC (topology 3) | Accuracy at the low-loss end. The surface-heated variant (topology 2) stays documented as the fallback. |
| Readout method | Early July | Flow readout, guard used only to null | Substitution keeps its role as the calibration method, not as the readout. |
| Preheat stage | Mid July | Adopted | RD-01's technique: pre-load the chamber to near the expected operating point with a known resistor before the DUT is switched on, which removes most of the first-point warm-up. The same resistor doubles as the calibration reference. |
| Gap skin finish | Mid July | Aluminium-foil lining, both facing skins | Radiation across the gap, not conduction, is the dominant leak with bare skins — see §5.4. Foil is a ~30× reduction and is what makes the guard spec affordable. |
| DUT-class modularity — copper cold block | Designed July; dropped 2026-08-27 | Removed from build scope | An optional water-cooled copper block was designed so a high-TDP DUT could dump heat straight into the water instead of into the chamber air. It was proven accuracy-neutral, but the F-series DUT dissipates about 15 W at its efficient operating points, where the block buys nothing. It was dropped as a time-box call, not a technical failure. See the note below. |
| Chamber temperature target and flow band | 2026-08-27 | Chamber Tss raised to 50–60 °C | Supervisor-driven. Raises the available ΔT and lowers the required flow. The consequences for the low-power corner are quantified in §7.9.2 and are not entirely benign. |
| Pump | 2026-08-27 | JwardTech 304K/BT peristaltic, stepper variant | Replaces the previously listed Kamoer KPHM600, which at 600–800 mL/min was about 10× too high-flow once the target ΔT rose. The stepper's low end goes cleanly toward zero, unlike a brushless pump's duty-cycle deadband. |
| Flow metering | 2026-08-27 | Biotech turbine sensor is the meter | Pump-as-flow-meter (step count × displacement) is not used for metrology; the pump is purely the flow actuator. |
| Chamber and wall sensing bus | 2026-08-27 | TMP117 over I²C for both chamber and guard array | Resolves an open three-way choice. Confirms I²C as the sensor bus, and therefore the need for an address multiplexer, since the guard array alone needs 8–12 sensors and the part has only four addresses. |
| Water ΔT sensing | 2026-08-27 | Jumo VIBROtemp pair, bare RTD | The ordered variant has no integrated transmitter, so a MAX31865-class front end per probe is required on the carrier board. |
| MCU and RTOS | August | ESP32-S3 DevKitC running Zephyr | Native Wi-Fi for live telemetry, existing Zephyr board support, and a module that plugs into a carrier board rather than a bespoke SoC design. |
| Supply rail | August | Single 12 V rail (LRS-350-12 class) | Sized bottom-up from the two 100 W heaters, fans and pump — about 22 A against the supply's 29 A. |

Table 8. Design decisions and revisions, in the order they were made.

On the copper block, and on notes that still mention it The rig was originally intended to carry an optional copper cold plate bolted to the DUT's baseplate, so that a high-TDP device could send most of its heat straight into the water (a split fraction α ≈ 0.97) rather than into the chamber air (α = 0). The analysis showed the reading is exactly independent of α — see §6.6 — so the block was never an accuracy feature; it was a way to widen the usable DUT range past ~200 W. Since the F-series DUT dissipates roughly 15 W at its efficient operating points and 200 W at worst, the block earns nothing on this DUT, and it was dropped from build scope on 2026-08-27 with the supervisor's agreement. The built rig is the α = 0, radiator-only case. Several working notes, some BOM links and the default alpha: 0.97 in params.yaml still refer to the block; they are historical and are not part of the delivered design.

# Thermal Model

## Modelling approach

The chamber is modelled as a lumped thermal network: a small number of bodies, each at one temperature, connected by conductances, each storing energy in a capacitance. This is the standard resistor-capacitor network of circuit theory with the labels changed, and the analogy is exact rather than decorative:

| Thermal | Electrical | Unit |
| --- | --- | --- |
| Temperature T | Voltage | K |
| Heat flow Q̇, power P | Current | W |
| Thermal conductance G, UA | Conductance | W/K |
| Thermal resistance Rth = 1/G | Resistance | K/W |
| Thermal capacitance C = m·c | Capacitance | J/K |
| First law at a node | Kirchhoff's current law | — |
| Time constant τ = C/G | τ = RC | s |

Table 9. The thermal-electrical analogy used throughout. The consequence worth stating: the whole model of §5.3 can be simulated unchanged in SPICE, with watts as current sources and kelvin above ambient as node voltages.

Three composition rules follow from the first law and are used everywhere. Paths that share the same temperature difference are in parallel and their conductances add; stages traversed in sequence are in series and their resistances add; a coolant stream entering at one temperature and leaving at the node's temperature behaves as a conductance to the inlet, Gadv = ṁ·cp.

All temperatures are written as deviations from ambient, θ = T − Ta, so ambient sits at zero and the ambient-drift disturbance appears explicitly as a −Ṫa term in every row rather than being scattered through the constants.

## Is lumping legitimate? — the Biot audit

A body may be represented by a single temperature when its internal conduction resistance is small compared with its surface film resistance, i.e. when the Biot number is small:

Bi = h·Lc / k = Rinternal / Rfilm ,   lump valid for Bi < 0.1		(4)

Every element inside the boundary was checked against this criterion, with deliberately high film coefficients so the test is conservative (fan-stirred air h = 50 W/m²K, water side 3 000, TIM contact 5 000).

| Element | Path checked | Bi | Verdict |
| --- | --- | --- | --- |
| Trim-heater copper foils | 35 µm Cu | 4 × 10⁻⁶ | Lump ✓ |
| Barrel resistor, ceramic core | Ø20 mm cylinder | 0.010 | Lump ✓ |
| Pickup-radiator fins | 0.1 mm fin | 2 × 10⁻⁵ | Lump ✓ |
| Radiator tube wall | 0.5 mm brass, water side | 0.014 | Lump ✓ |
| Wiring, Cu core | Ø1 mm | 3 × 10⁻⁵ | Lump ✓ |
| Fans, plastic impeller | ~2 mm PBT | 0.20 | Borderline — C ≈ 30–60 J/K, ~1 % of the total, and metered anyway |
| Sensor PCBs | 1.6 mm FR-4 | 0.13 | Borderline — C ≈ 2–5 J/K, negligible |
| Chamber and gap air | Fluid — Biot does not apply | — | Lumped by stirring: the fans' turnover time is seconds against minutes of system τ. This, not heat transfer, is the fans' real job — it is what makes Ti and Te mean something |
| RTD / TMP117 probes | Sheath, package | 0.15–0.17 | Not lumps — modelled as first-order sensor lags 1/(τs·s + 1) |
| XPS shells, 50 mm | Foam slab | ≈ 70 | Wall, not lump — by design. Bi ≫ 1 is the definition of a wall: the temperature drop lives inside the material, so it is modelled as a conductance (Ggap, Gout) |
| DUT board | 1.6 mm FR-4 into air or TIM | 0.27 … 27 | Fails — and that is why the instrument is a calorimeter. Junctions run tens of kelvin above the baseplate, so the DUT can never be one temperature. The reading never uses a DUT temperature: at dT/dt = 0 the internal gradients are constant and every dissipated watt must still leave along a metered path, however ugly the inside is |

Table 10. Biot audit. The two failures are structural features, not modelling errors: the foam is supposed to be a wall, and the DUT is supposed to be opaque to the measurement.

## The node equations

Three storage nodes survive the audit, plus one algebraic readout. With θ ≡ T − Ta throughout:

Ci·dθi/dt = (1−α)·PDUT + Paux − UArad·(θi − θw) − Ggap·(θi − θe) − Ci·Ṫa		(5a)

Cw·dθw/dt = α·PDUT + UArad·(θi − θw) − ṁcp·(θw − θin) − Cw·Ṫa		(5b)

Ce·dθe/dt = Pe + Ggap·(θi − θe) − Gout·θe − Ce·Ṫa		(5c)

Pmeas = ṁcp·(θw − θin) − Paux		(5d)

| Symbol | Meaning | Unit |
| --- | --- | --- |
| θi | Chamber air, above ambient | K |
| θw | Water side — resident loop water, radiator core, DUT mass | K |
| θe | Guard gap air | K |
| θin | Water inlet temperature from the reject loop | K |
| PDUT | The measurand — the DUT's dissipation | W |
| Paux | Metered auxiliaries inside B1 — chamber fans, preheat resistor | W |
| Pe | Guard heater — the null actuator | W |
| α | Fraction of the DUT's loss reaching the water directly rather than via the chamber air. α = 0 for the built rig | — |
| Ci, Cw, Ce | Capacitances of the three nodes | J/K |
| UArad | Chamber air ↔ water bridge through the pickup radiator | W/K |
| Ggap | Chamber → gap. The leak | W/K |
| Gout | Gap → ambient. Unmetered, and harmless | W/K |
| ṁcp | Advection through the water loop — the meter and the drain | W/K |
| Ṫa | Ambient drift rate | K/s |

Table 11. Symbols in the node equations.

Three derivatives, three states Only equations (5a)–(5c) carry a d/dt, so the model has exactly three states. Equation (5d) has none — it is a pure algebraic readout and belongs in the output matrix, never in the dynamics. This distinction is the one most often collapsed by mistake, and getting it wrong makes the reading appear to have dynamics it does not have.

One identity in (5b) and (5d) is worth pointing at, because it is why the instrument works at all: the water node's drain term and the meter's output are the same signal. ṁcp·(θw − θin) is simultaneously the plant's advection loss and the reported reading. That is also why θw must be a state — the reading is a capacitor voltage being observed, not an unknown being solved for.

## The conductances

Each conductance is a series/parallel assembly of the standard mechanisms.

Ggap = [ din/(λ·Awall) + 1/((hgap + hr)·Awall) ]⁻¹ + Gseam + Gft + Gwire		(6)

Gout = [ dout/(λ·Aout) + 1/((hext + hr,ext)·Aout) ]⁻¹		(7)

UArad = [ 1/(ηo·hair·Aair) + Rwall + 1/(hw·Awater) ]⁻¹		(8)

hr = 4·εeff·σ·T̄³ ,    εeff = (1/ε1 + 1/ε2 − 1)⁻¹		(9)

Equation (9) is the linearised radiation term, obtained by factoring T1⁴ − T2⁴ = (T1² + T2²)(T1 + T2)(T1 − T2) and freezing the bracket at 4T̄³. Over a 20–30 K band the error is a few percent of an already small term, and it is what keeps the whole model linear.

The foil rule — the single most load-bearing material choice Radiation across the gap runs skin-to-skin, in parallel with the foam's conduction, and it is the larger of the two if the skins are left bare. At T̄ = 320 K over the ≈0.6 m² gap: matte skins (ε ≈ 0.9 both sides) give εeff = 0.82, hr ≈ 6.1 W/m²K and Grad ≈ 3.7 W/K — 6.5× the foam's own 0.56 W/K. Aluminium foil on both facing skins (ε ≈ 0.05) gives εeff = 0.026 and Grad ≈ 0.11 W/K, about 30× less. This matters twice over. A perfect null cancels radiation along with everything else, but the null is never perfect, and the leak error is Ggap·δT with Ggap including Grad — so matte skins make every millikelvin of control error about seven times more expensive. Worse, the servo holds the gap air at Ti while radiation runs skin to skin, and the outer skin sits about 1 K below the gap air it feeds; with matte skins that alone leaks 3–4 W with a perfect air-side null, several times the entire budget. With foil, the same imperfection costs ~0.1 W.

| Quantity | Design value | Where it comes from |
| --- | --- | --- |
| Ggap | 0.56 W/K ⇒ Rth(in) = 1.79 K/W | XPS 50 mm, λ ≈ 0.035–0.04 W/m·K, A ≈ 0.6 m², foil-lined gap |
| Gout | 0.45 W/K | Outer shell, same class of construction |
| UArad | 30 W/K | PC-watercooling radiator with fans, air-side limited |
| ṁcp | 3–30 W/K design sweep (43–433 mL/min) | Set per operating point from ṁ = P/(cp·ΔTset) |
| hgap, hext | 5–15 W/m²K | Fan-stirred and free air respectively — the softest numbers in the model, which is why commissioning measures the assembled G instead |
| hr, foil-lined | ≈ 0.16 W/m²K | Equation (9) with εeff = 0.026 |

Table 12. Conductance design values. Commissioning values override all of these.

The ratio that makes the instrument work is UArad/Ggap ≈ 54. The metered path is about fifty times stiffer than the leak path, so the thermal current divider already sends the overwhelming majority of the heat into the water before the guard does anything; the guard then removes what the divider would still have lost. It is also why θi is pinned close to θw: a node tied to a rail through 30 W/K with only 0.56 W/K pulling elsewhere follows the rail.

## The capacitances

| Node | Contents | Design value | Expected range |
| --- | --- | --- | --- |
| Ci | Chamber air (rigid box ⇒ cv, not cp), participating foam skin, fixtures, looms | 650 J/K | 0.3–1.0 kJ/K |
| Cw | Resident loop water inside B1, radiator core, DUT mass (coupled through the baseplate) | 1 800 J/K | 1.0–2.7 kJ/K |
| Ce | Gap air, both facing foam skins, foil, guard heater | 300 J/K | 0.1–0.5 kJ/K |
| Ci + Cw | The mass the meter has to sweep out | 2 450 J/K | Sets the settling schedule and the steady gate |

Table 13. Capacitance budget. Only the wall's inner slice participates on the run timescale — the participating depth is δ ≈ 2·√(a·t), saturating at the full thickness once the run is long enough.

The foam wall's own distributed capacitance is real but is not one of the three states. It appears as an additional slow mode on top of the model, somewhere between 5.6 and 22 minutes depending on whether the slab is charged from one face or both — a factor of four that is still open and is the reason the first measurement point of a session takes hours while later points take minutes.

## The error budget

Relaxing both gates and keeping the deviations as first-order terms gives the error ledger. Each term is owned by a different part of the design:

δP = Ggap·δTnull + C·|dTi/dt| + cp,wΔTw·δṁ + ṁcp,w·δ(ΔTw) + δPaux + δPpar		(10)

| Term | At 15 W, ΔTw = 5 K | At 150 W, ΔTw = 10 K | Owner |
| --- | --- | --- | --- |
| Flow calibration, 0.5 % gravimetric | 0.08 W | 0.75 W | Calibration |
| ΔT pair, 20 mK matched after co-calibration | 0.06 W | 0.30 W | Sensing |
| cp,w property, 0.1 % | 0.02 W | 0.15 W | Property data |
| Auxiliary V·I metering, 0.5 % | 0.03 W | 0.05 W | Metering |
| Guard null residual, 0.3 K × Ggap (foil-lined) | 0.17 W | 0.17 W | Control loop |
| Storage, 0.2 W slice | 0.20 W | 0.20 W | Steady-state detector |
| Parasitics — anchored cabling, sealed seams | 0.05 W | 0.10 W | Build quality |
| RSS total against Pacc = ±1 W | ≈ 0.29 W ✓ | ≈ 0.87 W ✓ (tight) |  |

Table 14. Error budget at the two worst corners. The corners fail for different reasons: the cold corner is a null-plus-storage problem, owned by the control loops; the hot corner is a flow-calibration problem, owned by gravimetric rigour. Matte gap skins would put the null row at 1–4 W and sink both.

Two properties of this budget are worth stating plainly. First, the common-mode error of the ΔT pair cancels — only the pair mismatch after co-calibration survives, so absolute RTD accuracy is far less important than matching. Second, no control loop can fix a flow-calibration scale error: if the firmware computes with a nominal ṁcp while the loop drains at the true one, the reading is biased in proportion, and a 0.5 % flow error is 0.75 W at 150 W regardless of how well the guard is tuned.

# State-Space Model and Guard Control Design

The node equations of §5.3 are assembled into one state-space object. This is not a change of physics; it is the form that makes three questions answerable by inspection rather than by simulation: is the meter unbiased, how fast are the modes, and what does each error source cost in watts.

## States, inputs and output

x = [ θi , θw , θe ]ᵀ ,   u = [ Paux , Pe ]ᵀ ,   d = [ PDUT , θin , Ṫa ]ᵀ ,   y = Pmeas

The split between u and d is the control-theoretic one: u is what the firmware can command, d is what it must suffer. PDUT is a disturbance — in control terms, a measurand that cannot be set is a disturbance, however central it is to the purpose of the instrument.

ẋ = A·x + Bu·u + Bd·d ,    y = C·x + Du·u + Dd·d		(11)

## The matrices

Each row is one node's balance law divided by that node's own capacitance; each column is the state or input it is differentiated against. That is the whole derivation.

| A | θi | θw | θe |
| --- | --- | --- | --- |
| air | −(UArad + Ggap)/Ci | UArad/Ci | Ggap/Ci |
| water | UArad/Cw | −(UArad + ṁcp)/Cw | 0 |
| guard | Ggap/Ce | 0 | −(Ggap + Gout)/Ce |

Table 15. The system matrix A. Every entry has units of s⁻¹.

|  | Bu : Paux | Bu : Pe | Bd : PDUT | Bd : θin | Bd : Ṫa |
| --- | --- | --- | --- | --- | --- |
| air | 1/Ci | 0 | (1−α)/Ci | 0 | −1 |
| water | 0 | 0 | α/Cw | ṁcp/Cw | −1 |
| guard | 0 | 1/Ce | 0 | 0 | −1 |

Table 16. The input matrices. The Ṫa column is −1 in every row: that is the deviation coordinate system moving under the model, not heat flowing through a conductance.

C = [ 0 , ṁcp , 0 ] ,    Du = [ −1 , 0 ] ,    Dd = [ 0 , −ṁcp , 0 ]		(12)

The two non-zero D entries are the auxiliary-power subtraction and the inlet-temperature subtraction. They bypass the integrators entirely, which is why noise on Paux or on the inlet RTD appears in the reading instantly and unfiltered — a direct consequence of one matrix entry, and a reason to meter the auxiliaries well.

Note that ṁcp appears in four places — inside A, inside Bd, inside C and inside Dd. It is one conductance with two ends: the water node's drain, the rail it drains to, and the meter that watches it.

## Structural audit

Seven checks are run on A before any number is trusted. They catch wiring errors — signs, missing bridges, bad couplings — that no amount of simulation makes obvious. All seven are coded as assertions in calorimetry-ss.py: audit() and run on every plant build.

| # | Check | What it means | Catches |
| --- | --- | --- | --- |
| 1 | All diagonal entries negative | Every node decays on its own | A sign error that makes the simulation diverge |
| 2 | Row 1 sums to zero | The air has no exit except through the water and the guard | A missing path |
| 3 | Row 2 sums to −ṁcp/Cw | The water leaks to the inlet rail — this is the metered exit | A meter wired to the wrong rail |
| 4 | Row 3 sums to −Gout/Ce | The guard leaks to ambient — this is the unmetered exit, and the whole error term | A guard accidentally modelled as sealed |
| 5 | Ci·a12 = Cw·a21 = UArad; Ci·a13 = Ce·a31 = Ggap | Each shared bridge appears twice, with opposite signs, weighted by its own capacitance | Energy created at a coupling — the most common multi-node wiring error |
| 6 | a23 = a32 = 0 | Water and guard do not touch directly | A spurious coupling |
| 7 | All eigenvalues real and negative | A purely dissipative single-domain network cannot oscillate | A sign error anywhere. Only valid open-loop — see §6.7 |

Table 17. The structural audit.

Two non-zero row sums means two ways out of the boundary One of them is measured; the other is the error term. That single sentence is the entire physics of a calorimeter, and it is legible in the row sums of A before a single number is entered.

## Modes

At the 150 W / 10 K operating point (ṁcp = 15 W/K, 216 mL/min) the system matrix and its modes are:

A =  [ -0.047015   0.046154   0.000862 ]

[  0.016667  -0.025000   0.000000 ]   s^-1

[  0.001867   0.000000  -0.003367 ]

| Mode | τ | Eigenvector [θi, θw, θe] | What it physically is |
| --- | --- | --- | --- |
| Fast | 15.2 s | [+1.000, −0.408, −0.030] | Air and water exchanging across the radiator — opposite signs, guard uninvolved. In circuit terms, the two capacitors in series across the radiator bridge |
| Mid | 157.8 s | [+1.000, +0.893, −0.628] | Air and water moving together, draining through the water bus. The two capacitors in parallel, discharging through the advection conductance |
| Slow | 314.5 s | [+0.100, +0.077, +1.000] | The guard shell charging against ambient — almost entirely θe |

Table 18. Modes at ṁcp = 15 W/K. All three eigenvalues are real and negative, as they must be for a passive single-domain network.

The eigenvectors are the point of the exercise. Three time constants alone say "one fast thing, one medium thing, one slow thing"; the eigenvectors say which physical process each one is, and immediately that the five-minute mode is the guard shell rather than the measurement. Note also that the modal time constants are not the per-node C/ΣG figures (21 s, 40 s, 297 s): the fast mode is faster than either node alone and the mid mode is slower than both. Coupled nodes have modes, not individual time constants.

## The measurement, as a DC gain

The whole instrument reduces to one question: is the zero-frequency gain of the PDUT → Pmeas channel equal to one? With the guard open, it is not.

G(0) = −C·A⁻¹·B + D

| ṁcp [W/K] | Flow [mL/min] | DC gain, guard open | Deficit ε | Leakage-divider formula | Agreement |
| --- | --- | --- | --- | --- | --- |
| 3.00 | 43 | 0.916183 | 8.38 % | 0.083817 | 1.6 × 10⁻¹⁵ |
| 5.00 | 72 | 0.944985 | 5.50 % | 0.055015 | 9.2 × 10⁻¹⁶ |
| 7.50 | 108 | 0.960076 | 3.99 % | 0.039924 | 8.2 × 10⁻¹⁶ |
| 10.00 | 144 | 0.967804 | 3.22 % | 0.032196 | 3.8 × 10⁻¹⁶ |
| 15.00 | 216 | 0.975657 | 2.43 % | 0.024343 | 1.0 × 10⁻¹⁶ |
| 20.00 | 289 | 0.979631 | 2.04 % | 0.020369 | 2.6 × 10⁻¹⁶ |
| 25.00 | 361 | 0.982032 | 1.80 % | 0.017968 | 4.3 × 10⁻¹⁶ |
| 30.00 | 433 | 0.983638 | 1.64 % | 0.016362 | 1.0 × 10⁻¹⁶ |

Table 19. Systematic error with the guard open, α = 0. Computed two independent ways — from the state-space DC gain and from the closed form of equation (13) — and they agree to machine precision.

The deficit has a closed form. Writing everything as resistances R = 1/G:

ε = 1 − G(0) = [ Radv + (1−α)·Rrad ] / [ Radv + Rrad + Rgap + Rout ]		(13)

It is a leakage divider: the reading is short by the ratio of the metered path's resistance to the total resistance of all paths out of the boundary. Four consequences follow immediately:

| Limit | DC gain | Why |
| --- | --- | --- |
| Ggap → 0 — a perfect wall, or a guard that nulls θi − θe | Exactly 1 | No alternative exit exists |
| Gout → 0 — the guard sealed from ambient | Exactly 1 | Heat may enter the gap but cannot leave the boundary, so it must come back out through the water |
| Any α, with either of the above | Exactly 1 | The meter is physically incapable of telling how the heat reached the water |
| Small Ggap | 1 − Ggap·(Radv + (1−α)·Rrad) | The error is first order in the gap conductance — halve the leak, halve the error |

Table 20. Limits of the leakage divider.

The third row is the α-invariance result, and it is what made the copper block optional rather than necessary: the reading does not depend on how the heat got into the water. Verified numerically — at 150 W / 10 K the open-guard gain differs between the two configurations (0.9757 at α = 0, 0.9835 at α = 0.97) but the closed-loop gain is exactly 1.00000000 in both, with a residual error of order 10⁻¹² W, which is integrator convergence, not bias.

## Guard nulled — the exact meter response

Under the null gate (θe ≡ θi) the Ggap terms cancel on both sides and the plant collapses to a 2 × 2. Three DC gains then define the instrument:

| Channel | Value | Meaning |
| --- | --- | --- |
| PDUT → Pmeas | 1.000000 | Exact, for every α, Ci, Cw, UArad and ṁcp |
| θin → Pmeas | 0 | The reading is immune to the water setpoint — so auto-ranging the water ΔT between points is free |
| Paux → Pmeas | ≈ 0 (machine precision) | The Du = −1 subtraction provably works. Delete that one entry and every watt of preheat is reported as DUT loss |

Table 21. The three gains that define the instrument, with the guard nulled.

The transfer function of the meter, with the guard nulled, is second order with one zero:

Pmeas(s)/PDUT(s) = [ 1 + α·(Ci/UArad)·s ] / [ 1 + ((Ci+Cw)/ṁcp + Ci/UArad)·s + (Ci·Cw/(UArad·ṁcp))·s² ]		(14)

Its two poles separate cleanly because UArad ≫ ṁcp:

τfast ≈ (1/UArad)·(1/Ci + 1/Cw)⁻¹ ≈ 16 s ,    τslow ≈ (Ci + Cw)/ṁcp

τslow is the number that sets the schedule: it is the whole thermal mass inside the boundary being swept out through the meter. At 15 W and 43 mL/min that is 2 450/3 = 817 s, which is why the cold corner takes an hour to settle while the hot corner takes ten minutes. The zero sits at α·Ci/UArad — 21 s in the block configuration, and it disappears entirely at α = 0. The split changes the transient; it never changes the answer.

## The guard loop

Collecting the θe terms of (5c) gives the SISO plant the guard controller actually sees:

θe(s) = [ Ke/(τe·s + 1) ]·Pe(s) + [ β/(τe·s + 1) ]·θi(s)		(15)

Ke = 1/(Ggap + Gout) = 0.990 K/W ,   τe = Ce/(Ggap + Gout) = 297 s ,   β = Ggap/(Ggap + Gout) = 0.554

β = 0.554 is why the guard exists — and why proportional control is not enough Left unpowered, the gap settles at only 55 % of the chamber's temperature rise, because it loses the rest to the room through Gout. A pure proportional controller leaves a steady-state error of (1−β) divided by the loop gain, and that error is a permanent leak straight into the reading. The integrator is not a refinement — it is the accuracy. §7.8 prices it.

The controller is tuned by the λ-rule (internal model control), which sets one knob — the desired closed-loop time constant λg — and derives both gains from the identified plant:

Kp,g = τe/(Ke·λg) = 297/(0.990 × 60) = 5.000 W/K ,   Ki,g = Kp,g/τe = 0.01683 W/(K·s)		(16)

λg = 60 s is chosen as roughly a fifth of the open-loop τe — fast enough that the guard follows a moving chamber temperature, slow enough not to amplify sensor noise on a difference of two temperatures. Two nonlinearities sit on top and are mandatory:

•	Saturation, [0, Pmax]. The floor is zero, not −Pmax: heaters cannot cool. After a load dump the guard has no authority at all and the gap must bleed off through Gout on its own, which is slow. This is a real operational limit, visible in Figures 5 and 6.

•	Anti-windup, by back-calculation. The integrator is what makes the instrument exact, so the moment the heater clips, the thing that guarantees accuracy stops accumulating correctly unless it is explicitly unwound. The implementation adds (Pe − v)/Tt to the integrator rate, with Tt = Kp/Ki = τe, which is exactly zero whenever the output is not clipped.

The closed loop may oscillate — and that is not a fault The structural audit of Table 17 asserts that all eigenvalues are real, on the grounds that a purely dissipative single-domain network has nothing to trade energy with. That is true of the plant, and only of the plant. Once the guard PI is closed, the integrator is a second storage type — energy moves between the gap's thermal capacitance Ce and the controller's accumulated error, exactly as it does between L and C in a resonant circuit — and the closed loop legitimately acquires a complex pole pair. The audit must therefore never be run on the closed-loop matrix: complex eigenvalues in the plant mean a sign error; complex eigenvalues in the closed loop mean you have a controller. A damping ratio near 0.7 is well tuned; much below 0.4 says to slow λg down.

An optional feedforward term, Pff = Gout·(θtarget), can be added. It is free insurance: its DC gain to the reading's error is exactly zero, so a wrong Gout, a wrong ambient or a stale load estimate degrade the transient null error and nothing else. The integrator owns the steady null.

## What is deliberately not modelled

Two rules keep the model from growing without improving:

•	Anything whose datasheet or measured rating is the decision property. The radiator is a UA figure, not a fin geometry; the pump is a flow band and a steps-per-millilitre constant, not an impeller; the foam is a λ and a thickness, not a microstructure; radiation is a grey ε, not a spectrum. The justification is that commissioning measures the assembled C and G, so every sub-component detail is absorbed into two calibrated numbers — modelling it twice would add error, not fidelity.

•	Anything below the error floor. The standing promotion rule: a neglected effect earns a state or an error term only when its worst case exceeds about 10 % of Pacc, i.e. 0.1 W. Checked and dismissed on that basis: chamber-air cp versus cv, pump-head enthalpy, radiation through opaque foam, barometric drift, cp,w(T) beyond evaluating at the mean, and the DUT's internal thermal structure.

Four effects are genuinely nonlinear and are handled in the simulation rather than in the matrices: heater saturation, anti-windup, the bilinearity of ṁ (it multiplies a state, so the model is LTI only at a frozen flow — which is the operating protocol anyway), and sensor quantisation.

# System Simulation in Python

## Purpose and scope

The simulation exists to answer questions the algebra cannot: what the settling schedule actually is, how large the guard heater has to be, how big the transient null error gets before the loop catches up, and what happens when the actuator saturates. Saturation and anti-windup are not linear, so they cannot live inside A — which is why the model is integrated numerically rather than solved with a step-response function. Inside the limits it is exactly the LTI model of §6; outside them it is a different one.

## Files and how to run it

| File | Role |
| --- | --- |
| calorimetry-ss.py | The L1 plant. Builds A, Bu, Bd, C, Du, Dd at a frozen flow; runs the structural audit of Table 17 as assertions; provides unit helpers (mL/min ↔ kg/s ↔ W/K) and the modal decomposition. Runnable on its own — it prints the modes, the open-guard DC gains and a corner sweep. |
| calorimetry_sim.py | The main script. Wraps the guard PI, saturation and anti-windup around that plant, integrates the resulting 4-state system, rebuilds every probe signal from the trajectory, prints the probe map and writes the figures. |
| params.yaml | Every parameter, with its kind tag. Nothing numeric is hard-coded in either script. |

Table 22. The three source files, in ~/roboteq/python/.

$ cd ~/roboteq/python

$ python calorimetry_sim.py            # uses the defaults at the top of the file

$ python calorimetry_sim.py 300 12     # or: P_DUT = 300 W, water dT = 12 K

Dependencies are numpy, matplotlib, pyyaml and python-control. The script checks for all four against the interpreter that is actually running it and prints the exact pip line to fix a missing one, because the common failure here is having the packages installed against a different Python than the one being invoked. It also verifies that its output directory is writable before doing any work, and reports the size and wall-clock timestamp of every figure it writes — so that a figure that did not change can be distinguished from a viewer that is caching.

## The parameter file

| Parameter | Value | Kind | Note |
| --- | --- | --- | --- |
| C_i, C_w, C_e, C_r | 650, 1800, 300, 2000 J/K | estimate | Node capacitances; Cr is the reject-side reservoir |
| UA_rad | 30 W/K | estimate | Air ↔ water bridge |
| G_gap | 0.56 W/K | estimate | The leak |
| G_out | 0.45 W/K | estimate | Gap → ambient |
| alpha | 0.97 | design pick | Stale — the shipping configuration is 0.0. See the note in §4.7 |
| c_p_w, rho_w | 4180 J/kg·K, 995 kg/m³ | property | Water at ~305 K |
| mdot | 0.00358852 kg/s | estimate | Nominal only. Not a free parameter — the supervisor derives it per range as ṁ = P/(cp·ΔTset). This value is the 150 W / 10 K corner |
| ranges | 50 W → 5 K; 200 W → 10 K; 600 W → 25 K | control | The auto-ranging table |
| lambda_guard | 60 s | control | Guard closed-loop time constant |
| P_e_max | 100 W | design pick | Guard heater ceiling |
| P_acc | 1.0 W | control | Accuracy target |
| P_thr | 0.1 W | control | Residual threshold for the steady gate |
| sigma_pair_meter, sigma_pair_guard | 20 mK each | measurement | Sensor-pair residuals after co-calibration |
| flow_cal_err | 0.5 % | calibration | Gravimetric flow calibration |

Table 23. params.yaml. Kind tags mark whether a number is a design choice, a material property, a control target, a sensor spec or a value that commissioning will replace.

## Code structure

| Function | What it does |
| --- | --- |
| build_matrices(mcp) | Returns the six blocks at a frozen advection conductance. ṁ is bilinear — it multiplies a state — so the model is LTI only at a fixed flow, and a new plant is built per operating point. |
| audit(A, mcp) | The seven structural assertions of Table 17. Runs on every plant build. Documented as open-loop only: the closed loop legitimately grows a complex pair, so running this on Acl would fail for the wrong reason. |
| build_plant(mcp) | Wraps the matrices as a named python-control state-space object, and asserts that ṁcp is inside the design sweep — a units check that catches a flow entered in the wrong unit before it can produce plausible-looking nonsense. |
| modes(A) | Eigenvalues as time constants, slowest first, each with its normalised eigenvector — Table 18. |
| GuardPI | The controller: λ-tuned PI, output saturation with a floor of zero, and back-calculation anti-windup whose correction term is exactly zero when the output is not clipped. Gains are derived in __init__ from the plant constants, not typed in. |
| simulate(...) | Fixed-step RK4 over the 4-state system (3 plant states plus the guard integrator). Optionally pre-settles the rig with the DUT off for 30 000 s first, so t = 0 is a parked machine and the trace is a clean step response rather than a cold-start artefact. Returns a dictionary of named probe traces rebuilt from the stored trajectory. |
| demo_windup(...) | Steps the load, then dumps it, at four actuator ceilings. The ceilings are derived from the operating point rather than hard-coded, so the demonstration keeps its meaning when the power or ΔT is changed. |

Table 24. The main entry points.

## The probe map

Every signal the model exposes is rebuilt after integration and reported as a named probe, with its peak and its final value. The list is deliberately the same set of signals that the firmware will expose in telemetry, so that simulation traces and bench traces can be compared line for line.

| Probe | Expression | What it tells you | Unit |
| --- | --- | --- | --- |
| P_meas | ṁcp·(θw − θrail) − Paux | The reported reading | W |
| eps | Pmeas − PDUT | The error trace — the reason to build the model at all | W |
| e | θi − θe | The null error the guard controller acts on | K |
| P_leak | Ggap·e | The leak, in watts | W |
| P_e | Controller output after the limiter | Actuator demand against its ceiling | W |
| v | Controller output before the limiter | Saturation shows up as v ≠ Pe | W |
| P_rad | UArad·(θi − θw) | Watts crossing the pickup radiator | W |
| P_adv | ṁcp·(θw − θrail) | Watts carried by the flow | W |
| P_out | Gout·θe | Watts lost to ambient — unmetered | W |
| dT_w | θw − θrail | The auto-range variable | K |
| th_i, th_w, th_e | States | Node temperatures above ambient | K |
| x_I | Guard integrator state | Windup detector | K·s |
| sat_hi, sat_lo | Booleans | Pinned at the ceiling / at zero | — |

Table 25. The probe map.

## Simulation results

The figures below are the direct output of calorimetry_sim.py. All were run in the block-engaged configuration (α = 0.97), which is the script's default; as established in §6.5 this changes the transient shape but leaves the steady-state reading exactly unchanged, so the conclusions transfer to the built α = 0 rig. §7.7 gives the shipping-configuration numbers.

Each response figure has the same six panels:

| Panel | Shows | What to look for |
| --- | --- | --- |
| Top left — the measurement | PDUT (dashed, the truth) against Pmeas (solid, the reading) | That the reading converges onto the truth, and how long that takes |
| Top centre — eps = reading − truth | The error trace, with the ±1 W target marked. The y-axis is clipped because the t = 0 transient runs far off scale | The minute at which the error first enters the ±1 W band, and that it then goes to zero rather than to a small constant |
| Top right — node temperatures | θi (air), θw (water), θe (guard) and the inlet rail, all in kelvin above ambient | That θe tracks θi — the guard doing its job — and where the chamber settles relative to the rail |
| Bottom left — guard null error | e = θi − θe against the 1.79 K null gate | The transient overshoot, whether it breaches the gate, and how fast it returns |
| Bottom centre — actuator | Demand v (dashed) against delivered Pe (solid), with Pmax marked | Saturation, which shows as the two curves separating. Also the steady guard power, which is what sizes the heater |
| Bottom right — where the watts go | Prad (air→water), Padv (the reading), Pleak (air→gap), Pout (gap→ambient) | That Padv converges to the DUT power, that Pleak collapses to zero, and that Pout is non-zero but harmless |

Table 26. How to read the six-panel response figures.

### Nominal operating point — 150 W at 10 K water rise

Figure 1. Closed-loop step response at the nominal point: PDUT = 150 W, ΔTset = 10 K, ṁcp = 15 W/K (216 mL/min). The reading reaches 1 % of final in 12.9 min and enters the ±1 W band at about 14 min. The guard null error peaks at 1.76 K, just under the 1.79 K gate, and the leak peaks at 0.99 W before collapsing. Steady guard power is 6.8 W, against a 100 W ceiling — the actuator never comes close to saturating, and demand and delivered power are indistinguishable in the bottom-centre panel. In the bottom-right panel, note that Prad goes sharply negative in the first seconds: the water starts at the inlet rail temperature, below the chamber, so heat initially flows from water to air until the radiator bridge equalises. Note also that Pout settles to exactly the same value as the guard heater power — in steady state, with the leak nulled, the guard heater supplies precisely what the gap loses to the room, and none of it touches the reading.

### Same power, lower flow — the flow/settling trade

Running the same 150 W at a larger water ΔT means less flow, which is attractive for the ΔT-pair error budget but slow, because τslow = (Ci + Cw)/ṁcp. Figures 2 and 3 are the same load at 15 K and 20 K.

Figure 2. 150 W at ΔTset = 15 K, ṁcp = 10 W/K (144 mL/min). Settling stretches to 19.1 min and the null error peaks at 1.98 K, now above the 1.79 K gate. Steady guard power rises to 9.1 W because the chamber sits higher above ambient (θi ≈ 20 K rather than 15 K), so the gap loses more through Gout and the guard has to replace it.

Figure 3. 150 W at ΔTset = 20 K, ṁcp = 7.5 W/K (108 mL/min). Settling 25.4 min, null-error peak 2.12 K, steady guard power 11.3 W. Comparing Figures 1–3 gives the trade in one line: halving the flow doubles the settling time, raises the transient null error and raises the guard's steady power, while improving the ΔT-pair error by the same factor of two. The auto-ranging rule in params.yaml exists to place this trade per operating point rather than fixing it once.

### Top of the envelope — 200 W

Figure 4. 200 W at ΔTset = 8 K, ṁcp = 25 W/K (361 mL/min). The fastest corner in the set: 8.0 min to 1 %, inside the ±1 W band by about 9 min. Guard power is only 5.9 W steady, because the high flow pins the chamber lower (θi ≈ 13 K above ambient). The null error still peaks at 1.95 K, above the gate — the transient overshoot is driven by the mismatch between the fast chamber and the slow guard shell, not by the power level, and it is present at every corner.

The transient null error breaches the gate at and above ~100 W This is a genuine finding of the simulation and does not appear anywhere in the hand analysis, because it only exists once the guard is a real loop with a real τe instead of an assumed perfect null. The mechanism is a 20-to-1 speed mismatch: the chamber's fast mode is ~15 s while the guard shell's open-loop τe is 297 s, so the guard is chasing a reference that moves far faster than it can follow. It is a transient, so it does not bias a logged point — the null gate holds the reading until e is back inside 1.79 K. What it costs is schedule, and what it forbids is logging early. The candidate fix is to ramp the load and the feedforward together rather than stepping them, which is a sequencer change, not a model change.

### Actuator sizing, saturation and anti-windup

The last pair of figures answers a hardware question: how big does the guard heater have to be, and what happens if it is undersized. The load is stepped to the corner power, then dumped to 15 W partway through — the error reverses sign at that moment, which is precisely when a wound-up integrator shows itself. Four configurations are overlaid: the heater as designed (100 W), a marginal ceiling between the steady and peak demand with anti-windup on and off, and a deliberately undersized one below the steady demand.

Figure 5. Saturation and anti-windup at 150 W / 20 K, load dumped to 15 W at 25 min. Steady guard demand at this corner is 11.5 W with a transient peak of 16.6 W. Blue, 100 W as designed: never saturates, null error stays small, recovery after the dump takes 4.2 min. Orange and green, 13.3 W marginal: clipped 24 % of the run; the two traces are nearly indistinguishable, so with a ceiling this close to the demand, anti-windup is barely exercised (peak integrator 1 442 vs 1 473 K·s, recovery 5.2 vs 5.3 min). Red, 8.0 W undersized: clipped 48 % of the run, the integrator winds to 6 418 K·s — more than eight times the healthy case — the null error swings ±6 K, and recovery takes 22.4 min against 4.2. That is a sizing failure, and no amount of controller tuning fixes it.

Figure 6. The same study at 200 W / 8 K. Steady demand 5.9 W, transient peak 14.1 W — note that the steady demand is lower here than at 150 W / 20 K, because the high flow keeps the chamber cooler, while the transient peak is similar. An undersized 4.2 W heater is clipped for 58.7 % of the run and takes 17.4 min to recover. The design conclusion is that the guard heater must be sized against the transient peak, not the steady demand — roughly a factor of 2.5 apart at these corners. The specified 100 W barrel resistor gives about 6× margin over the worst transient peak seen, which also covers its second role as the guard-side preheat element.

Why every configuration has a slow tail after the load dump In all four traces the delivered power sits at exactly zero for several minutes after the dump, and the null error takes minutes to come back. That is not a controller problem: heaters cannot cool. Once the demand goes negative the output is pinned at the floor and the gap can only bleed its stored energy off through Gout, at its own τ. The consequence for the test sequence is that stepping the load down is slower than stepping it up, so a load map should be swept upward wherever possible.

## Corner sweep — the shipping configuration

The table below is the model run at α = 0, the built rig, across the design envelope. It is the single most useful output of the simulation, because it is what sets the test schedule.

| PDUT | ΔTset | ṁcp | Flow | Gain, guard open | Gain, guard closed | 1 % settling | Peak ¦e¦ | Peak leak | Pe steady |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 15 W | 5 K | 3.00 W/K | 43 mL/min | 0.9162 (−8.4 %) | 1.000000 | 63.6 min | 0.32 K ✓ | 0.18 W | 4.7 W |
| 50 W | 8 K | 6.25 W/K | 90 mL/min | 0.9540 (−4.6 %) | 1.000000 | 30.8 min | 1.07 K ✓ | 0.60 W | 6.6 W |
| 100 W | 10 K | 10.0 W/K | 144 mL/min | 0.9678 (−3.2 %) | 1.000000 | 19.4 min | 2.13 K ✗ | 1.19 W | 8.3 W |
| 150 W | 10 K | 15.0 W/K | 216 mL/min | 0.9757 (−2.4 %) | 1.000000 | 13.2 min | 3.16 K ✗ | 1.77 W | 9.0 W |
| 200 W | 10 K | 20.0 W/K | 289 mL/min | 0.9796 (−2.0 %) | 1.000000 | 10.1 min | 4.18 K ✗ | 2.34 W | 9.8 W |

Table 27. Closed-loop corner sweep at α = 0, flow matched to the load, step from a parked rig. ✓/✗ marks the 1.79 K null gate on the transient peak; every corner satisfies it in steady state, which is what the logged point depends on.

Three readings, two of which are warnings:

•	The closed-loop gain is exactly 1 at every corner. Five different flows, five different plants, one answer. This is the central result of the whole modelling effort: the instrument is unbiased by construction, not by calibration.

•	The cold corner is an hour. 63.6 min at 15 W. The mechanism is not the foam — it is that (Ci + Cw) ≈ 2.45 kJ/K has to be swept out through only 3 W/K of advection at 43 mL/min. Three levers exist, cheapest first: raise the flow and accept a smaller water ΔT; reduce Cw (every kJ/K removed is worth roughly five minutes); or estimate the steady value from the transient instead of waiting for it, which is a firmware change rather than a hardware one and is discussed in §7.9.2.

•	The transient null error grows with power and crosses the gate above about 100 W. As established above, it costs schedule rather than accuracy.

## Ablation — what the guard is actually worth

Same plant, same operating point, three guard configurations:

| Guard configuration | DC gain PDUT → P̂ | Deficit | Error at 150 W | Steady null error ess | Ggap·ess |
| --- | --- | --- | --- | --- | --- |
| PI — integral action | 1.000000 | 0.0000 % | 0.000 W | 0.00000 K | 0.0000 W |
| P-only (Ki = 0) | 0.994433 | 0.5567 % | 0.835 W | 1.49125 K | 0.8351 W |
| Guard OFF (open loop) | 0.967543 | 3.2457 % | 4.869 W | 8.69397 K | 4.8686 W |

Table 28. Guard ablation at 150 W / 10 K, α = 0. The identity ε = Ggap·ess holds exactly in the last two rows, which is a strong consistency check: the error in watts really is the null error in kelvin multiplied by the leak conductance.

Two conclusions, and the second is the less obvious one:

1.	The guard is load-bearing, not a refinement. Without it the rig is a 3.2 % instrument — 4.9 W at 150 W, five times outside the ±1 W target, and worse at low flow, where Table 19 shows the open-guard deficit reaching 8.4 %. This is the same failure mode Nair (RD-03) measured on a real unguarded build.

2.	What makes it exact is the integrator, not the hardware. Proportional control alone recovers 83 % of the deficit, from 4.87 W down to 0.84 W — but 0.84 W of the ±1 W budget is still spent on a single error term. The last 0.84 W is bought by one state, the integrator, and by nothing else. "Add integral action to the guard PI" and "make the calorimeter unbiased" are the same sentence — which is also why anti-windup is not optional.

## Findings and open items

### What the model settles

| # | Result |
| --- | --- |
| 1 | The reading is exactly unbiased with the guard closed — DC gain 1.000000 at every flow, every power and every α tested. |
| 2 | The systematic error with the guard open is a leakage divider, equation (13), verified against the state-space DC gain to machine precision across the flow range. |
| 3 | The reading is independent of α. The copper block was never an accuracy feature, which is what made dropping it a pure scope decision. |
| 4 | The guard's integral action is what makes the instrument exact — 3.25 % open, 0.56 % proportional-only, exactly zero with the integrator. |
| 5 | The water setpoint and the metered auxiliary power both have exactly zero DC gain to the error, so auto-ranging the water ΔT between points is free and the Paux subtraction is provably correct. |
| 6 | Guard heater sizing is set by the transient peak, not the steady demand — about 2.5× apart. The specified 100 W part has ~6× margin. |
| 7 | The settling schedule across the envelope: 10 min at 200 W to 64 min at 15 W, driven by (Ci + Cw)/ṁcp. |
| 8 | The steady gate can be derived rather than guessed: the slope threshold is the residual power threshold divided by the slowest observed time constant, computed per point. |

### Open items

| # | Item | Why it matters | Where it lands |
| --- | --- | --- | --- |
| 1 | The revised 30–40 K target and the flow floor. Treating the 50–60 °C chamber target directly as the water-loop ΔT puts the 15 W corner at 5–7 mL/min. Simulated, that corner has an open-guard gain of only 0.60–0.66, and 1 % settling of 6–8 hours. It also sits below the model's own sanity window (build_plant asserts 1 < ṁcp < 50 W/K, i.e. 14–721 mL/min). | The 30–40 K figure is a chamber-above-ambient target being used as a stand-in for the water-loop ΔT, and the two only converge if the pickup radiator is very effective. If it is applied literally to the water loop at low power, the cold corner becomes impractical. | Recommend keeping the auto-ranging table (5 K at 15 W, rising to 10 K by 200 W) and treating 50–60 °C as the chamber setpoint, not the water ΔT. Confirm against the first commissioning run. |
| 2 | Transient null error breaches the 1.79 K gate above ~100 W | Costs schedule and forbids early logging; new at closed loop | Candidate fix is a ramped load and feedforward instead of a step — a sequencer change |
| 3 | The foam wall's boundary condition — charged from one face or both | A factor of four (5.6 vs 22 min) on the first-point wait, which is on the critical path | A short Python job, not yet done |
| 4 | params.yaml still defaults to alpha: 0.97 | Harmless to the answer, misleading to a reader | One-line change |
| 5 | Model-based early stopping | Estimating the steady value from the transient collapses the cold-corner wait by roughly an order of magnitude, for about 40 lines of firmware and no hardware | Should be commissioned alongside the full wait for one campaign and compared, not shipped unvalidated |
| 6 | Monte-Carlo over the estimated parameter ranges | Every number in this report is a point estimate from mid-range design values | Turns each figure into a confidence interval |
| 7 | Sensor lags τs are placeholders (5 s) | They set the effective dead time the PI is tuned against | Commissioning |
| 8 | Zonal guard — the gap is modelled as one lump | A zero average null is not a zero leak; patchy faces still pay ΣGface·δTface | After single-zone tuning works |

Table 29. Open items, in priority order.

Commissioning overrides everything in this report Every conductance and capacitance above is a design estimate computed from geometry and material data. The commissioning step test — inject a known power with the calibration resistor, take C from the initial slope (C = P/Ṫ¦₀), G from the endpoint (G = P/(Tss − Ta)), and check that τ = C/G is consistent — measures the assembled rig and replaces them. The equations, the matrix structure, the audit and every conclusion in sections 5 to 7 survive that substitution unchanged; only the numbers move.

# Hardware Design

Section to be completed This section will cover the carrier PCB for the ESP32-S3 DevKitC: the 12 V → 5 V buck converter, the two independent heater power domains (inner chamber and guard) with their MOSFET stages and current sensing, the stepper driver for the peristaltic pump, PWM connectors for the fans, the RTD signal conditioning, and the protection strategy — ESD, overcurrent, overvoltage, short-circuit and thermal. It will also carry the bottom-up rail budget (≈22 A on the 12 V rail against a 29 A supply) and the measurement-grade versus protection-grade split on the current-sense channels.

# Firmware

Section to be completed This section will cover the Zephyr application on the ESP32-S3: the task split across the control, sensing, supervisor and telemetry threads; the devicetree description of the signal chains; the discretised guard loop (1 s sample rate, plant discretised zero-order-hold, controller Tustin) with its saturation and anti-windup; the null and steady gates of §4.6.3; the auto-ranging supervisor state machine; the staged power-up and heater interlocks; and the telemetry path — one JSON snapshot per second over WebSocket, rendered both by the local LVGL panel and by a browser dashboard.