Power-Converter Loss Calorimeter (LLC / Planar-Transformer DUT)

Thermodynamics Analysis

Governing equations and system transfer functions as the basis for configuring and simulating the calorimeter

| Author | Checked | Approved |
| --- | --- | --- |
| V. Gkionis | N. Adamopoulos | N. Adamopoulos |
| 2026-07-09 | Date | Date |

Revision History

| Iss. Rev. | Date | Author | Comments/Change log |
| --- | --- | --- | --- |
| 1.0 | 2026-07-09 | V. Gkionis | Initial release — thermodynamic modelling basis (P0): governing equations and system transfer functions. |

Table of Contents

Right-click this text and choose “Update Field” to build the table of contents.

# Purpose and scope

This document is the physics foundation (pipeline stage P0) for the calorimeter being developed to measure the power loss of Roboteq’s LLC resonant converter with planar transformer. Its single objective is to consolidate the thermodynamics of the whole instrument into one closed set of governing equations and transfer functions, so that every later stage — component selection, control-loop constant extraction, control design in Python/MATLAB, validation and assembly — inherits one calibrated model rather than a collection of rules of thumb.

The device under study is a water-loop, guard-nulled calorimeter (a double-jacket calorimeter, DJC). The device under test (DUT) is a switching power converter; for a 1.5 kW converter at efficiency η≥0.90 the dissipated loss is Ploss=(1−η)PDUT≤150 W, and the design must remain valid across the full working band of roughly 15–200 W. The instrument must resolve that loss to an accuracy budget of Pacc=±1 W.

The design goal is modularity: one permanent chamber, guard and loop (one form factor) that serves many DUT sizes and power ratings, with each new DUT handled by re-tuning only the operating point (pump flow and preheat target), not by rebuilding hardware. To make that possible every quantity below is written as a transparent function of a small set of design “knobs,” so the complete design can be recomputed — and simulated — for any choice of geometry, materials, or modules.

In scope: the governing energy balance, its derivation from first principles, the two lumped parameters (thermal capacitance C and conductance G) and their sub-models, wall losses as a function of wall thickness and chamber size, the water-loop hydraulics and thermal readout (including pressure), the air-circulation model, the guard/heater sizing, the steady-state and dynamic (transfer-function) behaviour, and the parametric recipe that ties them together.

Out of scope (deliberately): specific part numbers, vendor choices and bill-of-materials decisions. Those live in the materials and modules stage and are not treated here except where a physical constant (e.g. a radiator’s pressure rating) constrains an equation.

The working pipeline this note opens is:

physics (thermodynamics) → DUT specs (modularity window, and measuring other converters) → initial materials and modules pass → determine constants for the control loop → design the control loop (MATLAB or Python) → validate → assemble.

Everything that follows is the first arrow.

# The measurement principle

Calorimetry measures the total loss without resolving any single loss mechanism. Every loss channel inside the DUT — conduction, switching, magnetic core loss, gate drive, dielectric loss — is ordered electrical energy degrading irreversibly into heat. The second law of thermodynamics guarantees that degradation is one-way and complete, so whatever the mixture of mechanisms, it all arrives as the same joules of heat. One thermal measurement therefore captures the sum of all loss mechanisms at once, which is the fundamental reason calorimetry outperforms electrical input–output loss measurement for high-efficiency converters.

The instrument equation. Draw a control-volume boundary around the chamber interior. At steady state, with the wall leak nulled by the guard, all of the DUT’s dissipated power is carried away by the water, and the first law collapses to a single reading:

Ploss=m cp,w ΔT,  ΔT=Tout−Tin

where m is the coolant mass flow, cp,w its specific heat, and ΔT the temperature rise of the water across the chamber. Two co-calibrated resistance thermometers and one flow rate give the loss in watts. This is the entire instrument in one line.

Why the box must warm up first — the three destinations of every watt. Before steady state, the DUT’s power splits three ways, with no exceptions:

Ploss = m cp,w ΔT⏟into the water (measured) + Gwall (T−Ta)⏟wall leak (nulled) + i​Ci dTidt⏟storage (waited out)

The meter sees only the first term. The guard nulls the second. The third — storage — cannot be measured or nulled; it can only be waited out, and it dies exactly when nothing is changing temperature anymore, dTi/dt→0. That instant is steady state, and it is the moment the meter starts telling the truth.

Equilibrium versus steady state. The calorimeter never measures at equilibrium (nothing changes and nothing flows). It measures at steady state: temperatures hold still because the inflow from the DUT and the outflow through the water are balanced, with power streaming through the whole time. This distinction is what all “settling time” discussion is about.

# Governing equation — derivation from first principles

The dynamic behaviour of the chamber is captured by one first-order ordinary differential equation:

C dTdt=P−G (T−Ta)

This is not a new law. It is the first law of thermodynamics written for one lumped body, with its heat-loss term supplied by the transport rate laws. The derivation below makes it self-contained and reviewable.

## The three ingredients

First, energy is conserved: what accumulates inside a boundary equals what entered minus what left. Second, one modelling choice: treat the whole chamber interior as a single blob at one temperature T (the lumped assumption). Third, the rate laws: heat leaves a body in proportion to how far it sits above its surroundings.

## The six steps

Step 1 — Draw the boundary and give the blob one temperature. Enclose everything inside the chamber that warms together (air, inner wall skin, fixtures, DUT body) and assume it is close enough to one temperature to be described by a single number T(t). This is legitimate when heat moves around inside the blob much faster than it escapes — the small-Biot condition (Bi<0.1), quantified later.

Step 2 — Write the first law for the blob. The rate at which stored internal energy changes equals the rate in minus the rate out. The walls are rigid and nothing moves, so no mechanical work crosses the line; only dissipated power in and heat leaking out remain:

dUdt=Ein−Eout

Step 3 — Turn stored energy into temperature. For a solid or liquid, internal energy rises in proportion to temperature, with the heat capacity C=m cp as the constant of proportionality:

U−Uref=C (T−Tref) ⇒ dUdt=C dTdt

The left-hand side is now built; it is the exact thermal twin of a capacitor’s i=C dV/dt.

Step 4 — Name what goes in. Electrical power is dissipated inside the box (the preheat resistor, or the DUT’s own loss). The second law guarantees all of it degrades to heat, so

Ein=P

For the DUT, P is the net electrical power that stays inside as heat (power in through the wires minus any useful power carried back out). For preheat, P is the resistor’s V×I.

Step 5 — Name what leaves. Heat escapes to the room at Ta. Every escape path is linear in the same temperature difference T−Ta: conduction through the wall kAt(T−Ta), convection off surfaces hA (T−Ta), linearised radiation hrA (T−Ta), and the water loop carrying heat away mcp (T−Ta). Because every path shares the same driving difference, they combine into one conductance — parallel paths add conductances, series stages add resistances:

Eout=G (T−Ta),  G=1Rtot  [W/K]

G is where all the geometry and materials live: thicker walls, more foam, or a slower pump all reduce G.

Step 6 — Substitute. Putting steps 3–5 into step 2 gives the governing equation:

C dTdt=P−G (T−Ta)

Nothing else entered. The left side is energy piling up as a temperature rise; the right side is what is fed in minus what leaks out.

## Consequences read straight off the equation

Setting dT/dt=0 gives the steady state in one line, and integrating the linear ODE gives the full time response:

Tss=Ta+PG,    T(t)=Ta+PG1−e−t/τ,  τ=CG

Three facts fall out. The final value Tss=Ta+P/G is computable in advance from P and G — the steady state is predicted, not discovered by waiting, and this value is the preheat target. The wait is set by τ=C/G alone (63% at τ, 95% at 3τ, ~99% at 5τ); capacity sets the wait, never the destination. The measurement bias at any instant is exactly the storage term, so the meter under-reads by C dT/dt and reads Pmeter(t)=P (1−e−t/τ).

The equation is identical in form to a capacitor charging through a resistor, C dV/dt=I−V/R, which is why the whole model reasons with the RC analogy and why warm-up is characterised by the thermal time constant τ=C/G.

# System transfer functions

The governing equation is the seed; the instrument is a small network of such nodes, and its behaviour is fully described by a set of transfer functions. This section is the core modelling deliverable: it gives the input–output relationships that a control designer imports directly.

## Single node — the first-order lag

Take the deviation temperature θ=T−Ta and Laplace-transform the governing equation (θ(0)=0):

C s Θ(s)=Q(s)−G Θ(s) ⇒  Θ(s)Q(s)=1Cs+G=1/Gτs+1

Each thermal node is a first-order lag with DC gain 1/G (kelvin per watt) and time constant τ=C/G. The step response to a power step P is θ(t)=PG(1−e−t/τ), recovering the result above. This single pole at s=−1/τ is the building block of everything else.

## Multi-node — the state-space model

At design fidelity the chamber is 8–12 nodes: DUT surface, chamber air/contents, wall inner face, wall outer face / guard, coil metal, water inlet and outlet segments, and ambient as a fixed reference. Writing the first law per node and stacking gives a linear state-space system:

C T=−G T+q(t) ⇒ T=A T+B u,  A=−C−1G, B=C−1

Here C is the diagonal matrix of node capacitances Ci=micp,i, G is the conductance matrix (assembled by “stamping” each conductance between the nodes it connects, exactly like a circuit netlist), and q collects the injected powers (DUT loss, guard/preheat heaters, internal fans). The matrix transfer function from inputs to node temperatures is

H(s)=(sI−A)−1B=(sC+G)−1

Three standard results make this the single most useful object in the whole project:

- Steady state is the linear solve G T=q (set T=0) — instant wall gradients, leak and water ΔT for any candidate geometry, with no time-stepping.

- Dynamics / settling come from the eigenvalues λi of A: the time constants are τi=−1/λi, and the slowest τ is the settling time — obtained without simulating.

- Sensitivities for any design parameter p (wall thickness, dimensions, exchanger UA) are analytic: ∂T∂p=−G−1∂G∂p T — essentially free, and the basis of the geometry-optimisation decision gate.

The same (A,B) pair is exported directly as the plant for the control loops, so the sizing model and the control model are one model, never two.

## Worked illustration — a three-node reduction

A minimal but instructive reduction uses node 0 = DUT, node 1 = chamber air + internals, node 2 = wall (mid-thickness), with ambient fixed. With G01 from DUT to air, G12 from air to wall and G2a from wall to ambient:

C=C0000C1000C2,  G=G01−G010−G01G01+G12−G120−G12G12+G2a,  q=Ploss0G2aTa

Solving GT=q at Ploss=15 W with a purely passive wall floats the DUT node to roughly 57 °C, and the slowest eigenvalue gives a settling time of order 3 hours — which is precisely why a water loop is mandatory: adding the loop as one more conductance Gloop=mcp,w from the air node to a water-inlet reference collapses the DUT and air modes to minutes and slashes Tss to a survivable value. The passive result also proves, in five lines of algebra, that scaling Ploss to 200 W would drive an unguarded, un-cooled box past 400 °C.

## The measurement and guard transfer functions

Two input–output relationships matter operationally. The readout transfer function maps the loss to the water signal; at steady state it is the algebraic instrument equation Ploss=mcp,wΔT, and dynamically the water ΔT follows the air node through the pickup-exchanger lag. The guard transfer function maps guard-heater power to the inner–outer wall temperature difference δTwall=Ti−Te; the control objective is to drive δTwall→0 so that Gwall δTwall→0 and the leak term vanishes. Because the guard loop and the chamber/water loop are engineered to have time-scales separated by roughly 5–10×, they can be tuned as two nearly-independent single-input loops around the state-space plant above.

# Sub-model A: thermal capacitance C

The capacitance is the sum of mass times specific heat over everything inside the boundary that must warm:

Ctotal=i​mi cp,i=i​ρiVi cp,i

The air is thermally negligible (about 60 J/K for the whole interior); the real mass lives in the inner foam skin, the fixtures, the DUT body, and the water held in the loop, totalling on the order of Ctotal≈5000 J/K. Because each term is ρiVicp,i, changing the inner chamber volume changes C directly through the foam-skin and air terms — this is the term that both the chamber-sizing rule and the preheat-energy calculation act on.

Whether a part may be treated as one node is decided by the Biot number,

Bi=h Lk,  Bi<0.1⇒one node is legitimate.

The stirred water and the DUT heat-spreader pass this test; the foam wall does not (Bi≈4), so the wall is resolved as its own chain of nodes rather than lumped.

The capacitance sets two things that never affect the final reading but dominate the schedule: the time constant τ=C/G, and the preheat energy needed to bring the contents to their steady operating point,

Epre=i​Ci (Ti,ss−Ta)≈5000 J/K×20 K≈100 kJ.

# Sub-model B: the conductance network G

Every escape path is one of four transport rate laws, each linear in the same ΔT, so they assemble like a resistor network — parallel paths add conductances, series stages add resistances:

| Mechanism | Rate law | Conductance | Where in the rig |
| --- | --- | --- | --- |
| Conduction | Q=kAtΔT | Gwall=kAwallt | the foam wall (the leak) |
| Convection | Q=hA ΔT | Gconv=hA | air films on every surface |
| Radiation | Q=εσA(T14−T24) | Grad=hrA, hr≈4εσT3 | surface-to-surface IR |
| Advection | Q=mcp ΔT | Gloop=m cp,w | the water loop (the readout) |

These combine into the effective box-to-ambient conductance and its reciprocal resistance:

Gtotal=1Rtot=series stages​1parallel paths​G−1

G is where geometry and part choice enter the mathematics transparently. At 15 W the loop trickles and Gtotal≈2.5 W/K; at 150 W it runs hard and Gtotal≈8 W/K. Crucially the wall’s ≈0.56 W/K is dwarfed by the loop, so the loop, not the wall, is where the heat goes — which is exactly why the water circuit is the readout and the wall is only an error to be nulled.

## Wall losses as a function of thickness and chamber size

The wall leak is the parasitic term the whole guard exists to defeat, and it is a clean function of two knobs — wall thickness t and chamber size — through the conduction law:

Gwall=kfoam Awallt,  Awall=2 (LW+LH+WH)

with foam conductivity kfoam≈0.04 W/m·K. For the working size 0.45×0.45×0.25 m the wall area is Awall≈0.6 m2; at t=0.043 m this gives Gwall≈0.56 W/K, i.e. Rth≈1.8 K/W.

Two scaling laws govern the modularity trade-off:

- Thicker walls reduce the leak, but slow settling. Gwall∝1/t, so doubling foam thickness halves the leak conductance — but it adds foam mass and thus C, lengthening τ=C/G. Wall thickness is the classic conflicting objective (leak down versus settling up).

- Leak error scales as L2. Because Awall∝L2 for a fixed aspect ratio, the leak conductance and hence the residual leak error grow with the square of the linear size. Doubling the box quadruples Awall (to ≈2.2 W/K here), so the same guard null of ±0.3 K that costs only ±0.17 W at the working size would cost ±0.67 W — most of a ±1 W budget — on a box twice as large. This is the quantitative reason the sizing rule is “pack the parts, add one DUT-width of clearance, and stop.”

The residual leak the guard cannot remove is Gwall δTnull, where δTnull is the inner–outer temperature difference the guard leaves un-nulled (target ±0.3 K). This term goes straight into the error budget.

## Internal and radiative films

The air films on interior surfaces add resistances in series with the foam. Their coefficient h is the softest number in the whole model — a measured correlation rather than a first-principles quantity — with h≈3–10 W/m2K for natural convection and 10–50 when fan-forced. Radiation, linearised as hr≈4εσT3≈6ε W/m2K at chamber temperatures, is as strong as natural convection for bare matte surfaces (ε≈0.9) but is suppressed by an order of magnitude with a shiny foil lining (ε≈0.05). Because the foam’s 1.8 K/W dominates the films’ combined ≈0.3 K/W, the wall model can be kept simple: the films are a roughly 15% correction.

# The water loop — coolant circulation and hydraulics

The water loop is simultaneously the actuator that removes the heat and the meter that reads it. This section covers both the thermal readout and the hydraulic (pressure) behaviour.

## Thermal readout and flow sizing

The readout is advection, Ploss=mcp,wΔT, with m=ρwQv. The flow is designed so that ΔT lands in the 5–10 K signal band at the target power:

m=Pcp,w ΔT  ⇒  Qv=Pρw cp,w ΔT

The signal-to-error argument fixes the band: with a co-calibrated thermometer pair good to ±0.05 K, a ΔT of 5 K reads to 1% while a ΔT of 1 K reads to 5% — five times worse from the same probes. Hence ΔT is kept large (the signal), while the chamber-to-room difference is kept small (pure error). Sizing the flow at both power extremes gives the operating band that defines the modularity:

| Loss point | Design ΔT | Mass flow m | Volumetric flow Qv |
| --- | --- | --- | --- |
| 15 W (floor) | 5 K | 0.7×10−3 kg/s | ≈ 43 mL/min |
| 100 W | 5 K | 4.8×10−3 kg/s | ≈ 287 mL/min |
| 150 W (design) | 10 K | 3.6×10−3 kg/s | ≈ 215 mL/min |
| 200 W (ceiling) | 5 K | 9.5×10−3 kg/s | ≈ 573 mL/min |

Letting ΔT rise toward 10 K at the top end (which also improves the thermometer error budget) compresses the whole band to roughly 43–300 mL/min, which one positive-displacement pump on one tube size can cover. The power band therefore maps straight onto a pump-flow band — the core of the modular scheme, since each new DUT is served by choosing a flow setpoint, not by changing hardware.

## The two heat exchangers

The loop contains two exchangers: the pickup (air → water, inside the chamber) and the reject (water → room, outside). Each is a UA/effectiveness problem, sized from surface area and flow with either the log-mean-temperature-difference or the NTU-effectiveness form:

Q=UA ΔTlm ,    Q=ε Cmin ΔTmax , ε=f(NTU),  NTU=UACmin

with UApick,UArej≈15–25 W/K for the candidate radiators. These conductances are the series stages between the air node and the room in the network model.

## Hydraulics — flow regime, pressure drop and pump requirement

The loop’s pressure behaviour is set by the flow regime. The tube Reynolds number,

Re=ρwuDμ,  u=4QvπD2,

evaluates to roughly Re≈100–1500 across the whole flow band (for example Re≈870 at the 215 mL/min design point in an 8 mm bore), so the loop is laminar everywhere. Laminar flow has two consequences: radial mixing is poor, so static mixers must precede the thermometers; and the straight-tube pressure drop follows the Hagen–Poiseuille law,

Δp=128 μ L Qvπ D4 = 32 μ L uD2

which is strongly geometry-dependent through D−4. Worked at the design point (μ≈6.5×10−4 Pa·s at ~40 °C, L≈3 m, D=8 mm, Qv=3.58×10−6 m3/s) this gives Δp≈70 Pa — about 0.7 mbar, negligible. The corresponding pump head and hydraulic power are

H=Δpρwg≈7 mm,  Phyd=Δp Qv≈2.5×10−4 W.

The engineering consequence is that at these flows the straight tubing contributes almost nothing; the real loop resistance is dominated by the radiator fins and fittings, yet the total still sits far below the radiators’ 0.8 bar pressure rating in normal operation. The pressure risk is therefore not friction but dead-heading: a positive-displacement (peristaltic) pump driving into a closed valve will exceed 0.8 bar. The design rule that follows is to include a vented reservoir as a pressure relief and fill/bleed point, and to allow no valve that can fully close the loop. A positive-displacement pump also doubles as the flow meter (displacement per revolution × speed, calibrated gravimetrically), which is why the flow term in Ploss=mcp,wΔT can be known without a separate in-line sensor at these sub-PC-watercooling flows.

# Air circulation inside the chamber

The lumped capacitance and the single air temperature are only legitimate if the interior is well mixed. Left to itself the box air convects naturally at a Rayleigh number

Ra=g β ΔT L3ν α∼107,

which is messily transitional and would stratify the chamber. Two fans force convection instead, deliberately raising the internal film coefficient from natural values to hint≈10–50 W/m2K through a Nusselt-number correlation Nu=f(Re,Pr), with hint=Nu k/L. This forced film feeds two places in the network: the air-side film of the pickup exchanger, and the DUT-surface-to-air step.

One bookkeeping subtlety follows from putting fans inside the boundary: the fans dissipate their own electrical power inside the control volume, so at steady state the water carries Ploss+Pfan+Ppump,internal. Each auxiliary is therefore metered by its own V×I (DC, hence exact) and subtracted, or held constant so the calibration step removes it. Swapping a fan means re-metering its power; nothing else in the model moves.

# Guard and heating-element requirements

The guard is second-law judo: rather than fighting the outward wall leak, it removes the hill that drives it. A surface heater servos the wall’s outer face up to the inner temperature so that δTwall→0, making Gwall(T−Ta)→0 regardless of how large Gwall is. The heater is a servo sized for capability, not an open-loop source: the accuracy requirement is that the guard hold the wall gradient below

|Ti−Te|<Pacc Rth(in)=1.0 W×1.8 K/W=1.8 K.

## Total wattage

Held at the inner temperature, the guard supplies whatever the guarded surface loses to the room. For a room-facing copper-jacket build at the hot corner (Ti≈47 °C, Ta=25 °C):

Pheater≈(hout+hr) Ashell (Ti−Ta)≈(6+0.2)×0.86×22≈120 W,

so the supply is specified at roughly 1.5× that, i.e. about 150–200 W of capability on a 24 V rail. Bare copper helps twice: its low emissivity (ε≈0.04, so hr≈0.2) keeps the radiative term small, and the grounded copper skin doubles as an EMI screen.

## Resistance, drive and uniform flux

The heater is a copper foil whose resistance is built from geometry, since copper’s resistivity is tiny:

Rtape=ρe ℓtapewtape δtape,  ℓtape=Awallptape

For 10 mm-wide, 35 µm foil at 10 mm pitch over ≈0.8 m2 this is about ℓtape≈80 m and Rtape≈4 Ω. The power then follows from the rail through P=V2/R: 12 V gives only ~36 W (too little), while 24 V gives ~150 W at ~6 A — squarely in cheap-supply territory, and the same 24 V rail the pump uses.

Wiring the faces in series at a constant tape pitch makes the surface self-equalise. Every piece then carries the same current I, a piece of area A needs tape length ℓ=A/p so its resistance is R=ρeA/(p w δ)∝A, and its areal flux is independent of piece size:

q″=PpieceApiece=I2RA=I2 ρep w δ (independent of area),  Ppiece=q″ Apiece.

So a single current sets a uniform W/m² on every face automatically; per-face wattage is then just each face’s share of the area. Two construction rules protect the sensing: wind each face as a hairpin so adjacent turns carry opposite current and the ~80-turn solenoid field cancels, and drive the heater as smooth DC (linear or LC-filtered, never raw PWM) so it does not couple dB/dt into the thermometer leads.

## Preheat

Because the guard heater sits outside the nulled inner wall, its power never enters the loss reading, so it can also preheat the chamber to Tss before a run — pre-charging the thermal capacitance so the DUT’s own trickle need not spend hours filling it. The preheat target is the predicted Tss=Ta+P/G, and a crude electrical estimate of P sets it to within about 1 K, which is ample.

# Steady state, time constants and the two worst corners

Collecting the dynamic results: the destination is Tss=Ta+P/G, the wait is governed by τ=C/G, and the run is valid after about 5τ. Worked at the hot corner (150 W), the temperature ladder from a 25 °C room is:

| Rung | Conductance | Step at 150 W | Lands at |
| --- | --- | --- | --- |
| Room ← reject radiator (water inlet) | UArej≈15–25 W/K | ~6–10 K | water in ≈ 32–35 °C |
| Water heating through the box | mcp,w=15 W/K | 10 K by design | water out ≈ 42–45 °C |
| Air → water at the pickup radiator | UApick≈15–25 W/K | ~6–10 K | air ≈ 46–52 °C |
| DUT surface → air | DUT’s own hA | its own design | surface 60+ °C |

So “150 W” means the air parks near 50 °C and materials should be planned for ~55–60 °C. The two design corners are different problems:

- 150 W is the worst case for temperature — it sets material ratings, guard power and Tss. Here the loop runs fast (mcp,w=15 W/K, effective G≈8 W/K), so τ≈10 min and 5τ≈1 h.

- 15 W is the worst case for waiting — it sets settling time, preheat value and the per-point test schedule. Here the loop trickles (mcp,w≈3 W/K, effective G≈2.5 W/K), so τ≈33 min and 5τ≈3 h.

The design must satisfy both corners at once; the preheat scheme is aimed squarely at the slow low-power corner.

Detecting steady state with a number. Since the only un-nulled, un-measured term is storage, arrival is declared when the storage bias falls under its slice of the budget:

dTdt<PstoreCtotal=0.2 W5000 J/K=4×10−5 K/s≈0.14 K/h,

evaluated as a least-squares slope over a rolling 15–30 min window on every sensor, watching the slowest lump. This is the criterion that makes steady state detectable rather than guessed.

# Error budget and measurement validity

The instrument is trustworthy only when the two non-signal terms are controlled to inside the accuracy budget Pacc=±1 W. The two contributions are the residual wall leak Gwall δTnull (design ≈±0.17 W with a ±0.3 K null) and the storage bias Ctotal |dT/dt| (allocated 0.2 W, which sets the steady-state slope criterion above). The signal term itself is protected by keeping ΔT in the 5–10 K band and by co-calibrating the thermometer pair against each other, so that only their difference — not their absolute accuracy — enters the reading. The design passes when leak error plus storage bias stay within Pacc at both power corners.

# The parametric model — configuring and simulating any design

This section is the operational payoff: a single calculator that recomputes the entire design — and drives the simulation — for any choice of knobs. It is what makes the form factor modular rather than merely documented.

## The master ledger

Every variable and constant, grouped. Legend for Kind: knob = chosen (a modular input); prop = fixed material/physical constant; derived = computed; target = a spec/budget; meas = read from a sensor at run time.

Group A — the measurand

| Symbol | Meaning | Units | Value / band | Kind |
| --- | --- | --- | --- | --- |
| PDUT | DUT input electrical power | W | up to 1.5 kW | knob |
| η | DUT efficiency | — | ≥ 0.90 | knob |
| Ploss | dissipated loss (the measurand) | W | (1−η)PDUT; ~15–200 | derived/target |

Group B — geometry (modular knobs)

| Symbol | Meaning | Units | Value / band | Kind |
| --- | --- | --- | --- | --- |
| L,W,H | inner chamber dimensions | m | ≈ 0.45 × 0.45 × 0.25 | knob |
| V | inner free volume | m³ | 0.03–0.06 | knob |
| Awall | inner wall area, 2(LW+LH+WH) | m² | ≈ 0.6 | derived |
| t | wall (foam) thickness | m | 0.03–0.05 | knob |
| ptape,wtape,δtape | heater-tape pitch, width, foil thickness | m | 0.010, 0.010, 35 µm | knob |
| ℓtape | total tape length, Awall/ptape | m | ≈ 80 | derived |
| D | coolant tube inner bore | m | tubing choice | knob |
| Apick,Arej | pickup / reject radiator area | m² | radiator choice | knob |

Group C — material and physical properties (constants)

| Symbol | Meaning | Units | Value | Kind |
| --- | --- | --- | --- | --- |
| kfoam | foam thermal conductivity | W/m·K | ≈ 0.04 | prop |
| cp,w | specific heat of water | J/kg·K | ≈ 4180 | prop |
| ρw | water density | kg/m³ | ≈ 1000 | prop |
| μ | water dynamic viscosity | Pa·s | ≈ 0.65–1.0 ×10⁻³ | prop |
| ρe | copper electrical resistivity | Ω·m | 1.68×10⁻⁸ | prop |
| ε | surface emissivity | — | foil ~0.04; matte ~0.9 | prop/knob |
| σ | Stefan–Boltzmann constant | W/m²·K⁴ | 5.67×10⁻⁸ | prop |
| βair | air expansion coeff, ≈1/T | 1/K | ≈ 0.0033 | prop |
| Pr | Prandtl number | — | air 0.7, water ~5 | prop |

Group D — the two lumped parameters (the heart of the model)

| Symbol | Meaning | Units | Value / band | Kind |
| --- | --- | --- | --- | --- |
| Ctotal | total heat capacity | J/K | ≈ 5000 | derived |
| Gwall | wall leak, kAwall/t | W/K | ≈ 0.56 | derived |
| Gloop | coolant advective conductance, mcp,w | W/K | 3–40 | derived |
| Gtotal | effective box → away | W/K | ≈ 2.5 @15 W … 8 @150 W | derived |
| Rtot | total thermal resistance, 1/Gtotal | K/W | — | derived |
| τ | time constant, Ctotal/Gtotal | s | ≈ 33 min @15 W, 10 min @150 W | derived |

Group E — coolant circulation

| Symbol | Meaning | Units | Value / band | Kind |
| --- | --- | --- | --- | --- |
| m | coolant mass flow, ρwQv | kg/s | 7×10⁻⁴ … 9.5×10⁻³ | knob (pump) |
| Qv | volumetric flow | mL/min | 43–573 (≈215 @150 W) | knob |
| ΔT | coolant rise Tout−Tin (the signal) | K | 5–10 by design | target |
| Δp | loop pressure drop, 128μLQv/(πD4) | Pa | ≪ 0.8 bar in normal flow | derived |
| UApick,UArej | exchanger conductances | W/K | ≈ 15–25 | knob/derived |

Group F — air circulation

| Symbol | Meaning | Units | Value / band | Kind |
| --- | --- | --- | --- | --- |
| Nfan,dfan | fan count, diameter | —, mm | 2 × 140 | knob |
| hint | internal forced-convection coeff | W/m²·K | 10–50 | derived |
| hr | radiation film, 4εσT3 | W/m²·K | ≈ 6ε | derived |
| Re | tube Reynolds, ρuD/μ | — | 100–1500 (laminar) | derived |
| Ra | box Rayleigh | — | ~10⁷ | derived |
| Bi | Biot, hL/k | — | foam ~4; contents <0.1 | derived |

Group G — preheat and guard

| Symbol | Meaning | Units | Value / band | Kind |
| --- | --- | --- | --- | --- |
| Pheater | guard / preheat power | W | 50–150 (spec ~150–200) | knob/derived |
| Rtape | heater resistance, ρeℓtape/(wtapeδtape) | Ω | ≈ 4 | derived |
| Vrail,Ih | heater rail voltage, current | V, A | 24, ~6 | knob |
| δTnull | guard residual across the wall | K | ±0.3 | target |
| Epre | preheat energy, i​Ci(Ti,ss−Ta) | J | ≈ 100 k | derived |

Group H — targets and budgets (the gates)

| Symbol | Meaning | Units | Value | Kind |
| --- | --- | --- | --- | --- |
| Pacc | accuracy budget | W | ±1 | target |
| Pstore | budget slice for storage bias | W | 0.2 | target |
| ¦dT/dt¦max | steady-state criterion, Pstore/Ctotal | K/s | ≈ 4×10⁻⁵ (0.14 K/h) | derived |
| tsettle | settling time gate | s | ≤ ~2 h | target |
| Ta | room temperature | K | ≈ 298 | knob |
| Tss | steady-state temperature, Ta+P/G | K | ≈ 320 (≈47 °C) @150 W | derived |
| Tmax | material limit | °C | plan 55–60 | target |

## The compute order

Given the knobs, the design is recomputed in this order — each step is one of the equations above:

- Geometry → Awall=2(LW+LH+WH), V, ℓtape=Awall/ptape.

- Capacitance → Ctotal=i​ρiVicp,i.

- Conductances → Gwall=kAwall/t; Gloop=mcp,w; assemble Gtotal(m) and Rtot.

- Dynamics and steady state, evaluated at both corners (15 W and 150 W) → τ=C/G; Tss=Ta+P/G; ΔT=P/(mcp,w).

- Actuators and errors → Pheater=hAwall(Ts−Ta); leak error =Gwall δTnull; storage bias =Ctotal|dT/dt|; loop pressure Δp=128μLQv/(πD4).

## The design gates

Iterate the knobs until both corners clear every gate, then freeze the geometry — the frozen set is the permanent form factor:

- Is ΔT∈[5,10] K at the design flow? If not, retune m.

- Is 5τ≤tsettle (~2 h) at 15 W? If not, shrink C (smaller box / less foam mass), raise G (stronger pickup), or lean harder on preheat.

- Are Tss and Tmax within material limits at 150 W?

- Is leak error + storage bias ≤Pacc?

- Is Δp safely below the 0.8 bar loop rating, with a vented reservoir and no closable valve?

Commissioning then calibrates the frozen design once: a single step-response test with a known calibration power Pcal yields all three constants directly — C=Pcal/(dT/dt)|t=0 (initial slope, nothing leaking yet), G=Pcal/(Tss−Ta) (final rise, nothing storing anymore), and τ from the exponential fit, with the consistency check τ≈C/G. After that the model is predictive for any power in the band, and each new DUT is a two-number change (flow setpoint and preheat target), not a rebuild.

# Simulating the model

The same equations are the simulation. The recommended stack solves the model at the lowest fidelity rung that answers each question and climbs only on evidence.

At the design rung the chamber is the linear system CT=−GT+q. Steady state is one linear solve GT=q; settling comes from the eigenvalues of A=−C−1G; sensitivities are analytic. In Python this is a few lines of numpy/scipy (numpy.linalg.solve for steady state, scipy.linalg.eig for time constants, scipy.integrate.solve_ivp for transients); the same (A,B) exports to MATLAB/Simulink as a state-space block for the control loops, so one model serves sizing, optimisation and control.

Two numerical cautions. The rig’s time constants span seconds (thermometer tip) to hours (foam box), a stiffness ratio of 103–104, so an implicit integrator (BDF/Radau in scipy, ode15s in MATLAB) is mandatory; an explicit solver such as ode45 will crawl. Where the wall gradient is resolved as a field, the explicit finite-difference update is stable only for a Fourier number Fo=α Δt/Δx2≤12; implicit schemes are unconditionally stable.

Every simulation earns trust in stages: analytic checks (lumped cooling must give e−t/RC; a steady wall must give kΔT/t), grid/step convergence, a steady-state conservation audit (power in equals power out across every boundary), and finally the P6 step-response test that replaces estimated C,G,τ with as-built values.

# References

Thermodynamic laws (the frame). First law / energy conservation; second law / entropy; zeroth law / temperature.

Transport and modelling (the C and G sub-models). Heat capacity; Fourier’s law / conduction; Newton’s law of cooling / convection; lumped-element model and Biot number; Stefan–Boltzmann radiation; Nusselt, Rayleigh and Prandtl numbers; Hagen–Poiseuille flow; heat-exchanger sizing by NTU-effectiveness and LMTD; CoolProp for fluid properties.

Papers that use this method our way. Badstübner et al., “Calorimetric Power Loss Measurement for Highly Efficient Converters,” ETH Zürich PES / IPEC 2010 — the lumped energy-balance plus preheat approach (IEEE Xplore). Simpson & Hopkins, “An Accurate and Flexible Calorimeter Topology…,” IEMDC 2017 (Bristol) — the surface-heated guard that is the wall-nulling actuator (DOI). Kosonen et al., “Calorimetric concept for measurement of power losses up to 2 kW in electric drives,” IET EPA 2013 — the precedent for one rig serving many ratings (IET).